[image: image1.png]


 

                   O2 Sensors
	COMPANY OPERATING PROCEDURES


Teledyne R21A (PN 0110121), 
O2 Sensor Test Procedure
VM3/COP/37.36
	Date:
  3-Feb-03
	Revision date: 23-Aug-18
	Issue:   3



Test & Equipment Required

Functional test required using R21 tester and R21 test adapter, (digital multimeter and adapter).

Output specification as in document: (VM3COP37.00, column (i)).

Method.

1. Open cell packaging and remove O2 sensor.

2. Check the O2 sensor for damage and signs of leaking electrolyte. 

3. Connect the O2 sensor to the R21 tester, select 200mV range (marked with arrows) and wait a few seconds for meter reading to stabilize.


4. Check for an output  Ensure that meter reading lies within the output specification range,
9mV – 14mV

5. Scan into Intrastats as Passed with output in the QA Field

6. If no, or very low output, or sensor is returned as faulty leave the O2 sensor exposed to air for 1 Hour Minimum.


7. If the O2 sensor fails the output specification test, it should be left sealed and  placed in the ducket on the QA bench labeled Return to Supplier. Each sensor should have its barcode attached with the output / problem on it. Intrastats should be filled in as a Fail with the output logged into intrastats.

8. Disconnect the O2 sensor from the R21 tester and label with Viamed stickers if appropriate.

9. Reseal in serial numbered packet, ibox and add viamed outer box label.

VM3COP37.36 R-21A Oxygen Sensor 0110121_OUT_OF_DATE_12_Jul_2018.doc    23/08/2018         Page 1 of 1

[image: image1.png]