[image: image3.jpg]




Specification Datasheet - 8010012

Innerspace Systems Corporation

SC-I22D Oxygen Sensor
[image: image4.jpg]



v1.5
11th July 2016
	Specifications


	VST part number
	8010012

	OEM model number
	SC-I22D

	Output
	9 – 13 mV

	Output connector
	3 pin Molex

	Input
	Threaded M16 x 1mm

	Input O ring
	Yes

	Range
	0 – 100% oxygen

	Pressure range
	600 to 2500 mbar

	Repeatability
	<1% volume oxygen at constant temperature and pressure

	Linearity error
	0.1 to 1.0 bar < 3% relative

0.1 to 2.5 bar < 5% relative 

(at ppO2 > 1.0 bar after 10 min equilibration time)

	Warm up time
	< 30 minutes after sensor installed 

	Zero offset voltage
	< 0.5% oxygen reading in 100% nitrogen @ 25ºC after 36 seconds

	Cross interference
	<0.5% volume oxygen response to 5% CO2 balance N2

	Response time
	< 5 s for 90% of final value.

	Operating humidity
	0 – 99% R.H. (non-condensing).

	Influence of humidity
	- 0.03% relative per % R.H. at 25 ºC.

	Influence of mechanical shock
	< 1% after fall from 1m

	Operating temperature range
	0 – 50 ºC

	Storage temperature
	-20 to +50 ºC

	Recommended storage temperature
	+5 to +15 ºC

	Temperature compensation
	Integrated NTC network

	Temperature compensation error
	0 °C to 50 °C, 5% relative error

	Nominal sensor life
	> 500,000 % oxygen hours

	Long term output drift
	< 1% volume oxygen per month, typically < -15% over sensor lifetime.

	Shelf life
	24 months

	Load required
	≥ 10 K Ohms

	Input - membrane protection grille
	No

	Input - hydrophobic membrane
	Yes

	Conformal coating on PCB
	Yes

	Pressure equalization - PCB
	Yes 

	Pressure equalization - output cap
	None

	Weight
	25 g

	Flow diverter
	No

	Packaging
	Gas barrier bag

	Shipment packaging
	Bulk packed

	Warranty period
	12 months from date of sales invoice

	All specifications are applicable at standard conditions: 1013 mbar, 25°C dry ambient air.



Page 2 of 4

	Drawings/images


[image: image5.jpg]35
416

215
7
||
M16x1
(5/8-24 UNEF)
@28

Locking Latch

) ()

|
I

3 PIN MOLEX CONNECTOR




[image: image6.jpg]



[image: image7.jpg]



[image: image8.png]Vandagraph
>ensor Technologies





Page 3 of 4

	Labelling


Oxygen sensor label 90.0mm x 15.0mm

White vinyl label, text & graphics – black

[image: image1.jpg]N

% Innerspace Systems Corporation ~ SC-122D Oxygen Sensor

171 Midway Court ™M SN 1SCxoooxc
entralia,
AN % CE Lo i




Gas barrier bag label 63.5mm x 46.6mm

White label, text & graphics – black

[image: image2.jpg]g,
-W"“ Innerspace Systems Corporation

info@customrebreathers.com

Emai
Phone: (+001) 360.330.9018

1711 Midway Court
Centralia, WA 98531

USA

SC-122D Oxygen Sensor

M 2011-07 [:E] C€
W *®

This oxygen sensor contains caustic liquid, in case of leakage avoid contact with eyes
and skin. Do not puncture or staple this bag. Do not remove the oxygen sensor from
this bag until ready to use. Depending upon the ambient conditions the oxygen sensor

may require 30 minutes to reach signal stability after installation.





Page 4 of 4
Vandagraph Sensor Technologies Ltd.

15 Station Road, Cross Hills, Keighley, West Yorkshire BD20 7DT, United Kingdom

Tel +44 (0) 1535 634900   Fax +44 (0) 1535 635582   e-mail: info@vandagraphst.com

www.vandagraphst.com

