[image: image3.jpg]




OEM Specification Datasheet - 8010004

JJ-CCR

R17JJ-CCR Oxygen Sensor
[image: image4.jpg]



© Vandagraph Sensor Technologies Ltd.

v1.3
11th July 2016
	Specifications


	VST part number
	8010004

	OEM model number
	R17JJ-CCR Oxygen Sensor

	Output
	9 – 13 mV

	Output connector
	Radiall SMB Series R114.186.000

Right angled plug, crimp type

	Cable
	Belden RG174AU, length 15 cm

	Input
	Threaded M16 x 1mm

	Input O ring
	No

	Range
	0 – 100% Oxygen

	Pressure range
	600 to 2500 mbar

	Repeatability
	<1% volume oxygen at constant temperature and pressure

	Linearity error
	0.1 to 1.0 bar < 3% relative

0.1 to 2.5 bar < 5% relative

(at ppO2 > 1.0 bar after 1 hour equilibration time)

	Warm up time
	< 30 minutes after sensor installed 

	Zero offset voltage
	< 0.5% oxygen reading in 100% nitrogen @ 25ºC after 36 seconds

	Cross interference
	<0.5% vol. O2 response to 5% CO2 balance N2

	Response time
	< 5 s for 90% of final value.

	Operating humidity
	0 – 99% R.H. (non-condensing).

	Influence of humidity
	- 0.03% relative per % R.H. at 25 ºC.

	Influence of mechanical shock
	< 1% after fall from 1m

	Operating temperature range
	0 – 50 ºC

	Storage temperature
	-20 to +50 ºC

	Recommended storage temperature
	+5 to +15 ºC

	Temperature compensation
	Integrated NTC network

	Temperature compensation error
	0 °C to 50 °C, 5% relative error

	Nominal sensor life
	> 500,000 % oxygen hours

	Long term output drift
	< 1% volume oxygen per month, typically < -15% over sensor lifetime.

	Shelf life
	24 months

	Load required
	> 10 K Ohms

	Input - membrane protection grille
	No

	Input - hydrophobic membrane
	Yes 

	Conformal coating on PCB
	Yes

	Pressure equalization - PCB
	Yes – 2 x 3mm holes

	Pressure equalization - output cap
	Yes – 1 x 2mm hole

	Weight
	31 g

	Flow diverter
	No

	Packaging
	Gas barrier bag - transparent

	Warranty period
	12 months from date of sales invoice

	All specifications are applicable at standard conditions: 1013 mbar, 25°C dry ambient air.



Specifications subject to change 

Page 2 of 4

	Drawings/images


[image: image5.jpg]


[image: image6.jpg]160%5

345
a1

M16x1 (5/8-24 UNEF)

$29+0.2





[image: image7.jpg]



[image: image8.png]Vandagraph
>ensor Technologies









All dimensions are in mm

Page 3 of 4

	Labelling


Oxygen sensor label 

90.0mm x 15.0mm

White vinyl label, text & graphics - black

[image: image1.png]JJ CCR Aps R17JJ-CCR Oxygen Sensor
SIN  JJ123456
e o penmanc 1\ | 98] CE zm N T T





Gas barrier bag label 

63.5mm x 46.6mm

White label, text & graphics – black

[image: image2.png]JJ-CCR ApS

Dyrlevvej 11
DK-4720 Presto
Denmark

R17JJ-CCR Oxygen Sensor

(M 2010-11

SIN  JJ123456

This oxygen sensor contains caustic liquid, in case of leakage avoid contact with eyes
and skin. Do not puncture or staple this bag. Do not remove the oxygen sensor from
this bag until ready to use. Depending upon the ambient conditions the oxygen sensor
may require 30 minutes to reach signal stability after installation.





NOTE: Serial number prefix to be  JJ e.g. JJ100001
Page 4 of 4
+ve           -ve








Vandagraph Sensor Technologies

15 Station Road, Cross Hills, Keighley, West Yorkshire BD20 7DT, United Kingdom

Tel +44 (0) 1535 634900   Fax +44 (0) 1535 635582   e-mail: info@vandagraphst.com

