[image: image1.jpg]




Specification Datasheet - 8010016

Juergensen Marine

HH22 Hammerhead Oxygen Sensor
[image: image2.png]W Juergensen Marine HH22 Hammerhead Oxygen Sensor

pooen 7 P

Distributed by: Silent Thrill Limited

Distibued by @ SIN JM1123456
44 10y 208 123 6066 A g C€ 2013-03 Il ININKNRHCAI(ACERA





v1.4
19th March 2018
	Specifications


	VST part number
	8010016

	OEM
	Juergensen Marine

	OEM model number
	HH22 Hammerhead Oxygen Sensor

	Output
	9 – 13 mV

	Output connector
	3 pin Molex

	Input
	Threaded M16 x 1mm

	Input O ring
	Yes

	Range
	0 – 100% oxygen

	Pressure range
	600 to 2500 mbar

	Repeatability
	<1% volume oxygen at constant temperature and pressure

	Linearity error
	0.1 to 1.0 bar < 3% relative

0.1 to 2.5 bar < 5% relative

(at ppO2 > 1.0 bar after 1 hour equilibration time)

	Warm up time
	< 30 minutes after sensor installed 

	Zero offset voltage
	< 0.5% oxygen reading in 100% nitrogen @ 25ºC after 36 seconds

	Cross interference
	<0.5% volume oxygen response to 5% CO2 balance N2

	Response time
	< 5 s for 90% of final value.

	Operating humidity
	0 – 99% R.H. (non-condensing).

	Influence of humidity
	- 0.03% relative per % R.H. at 25 ºC.

	Influence of mechanical shock
	< 1% after fall from 1m

	Operating temperature range
	0 – 50 ºC

	Storage temperature
	-20 to +50 ºC

	Recommended storage temperature
	+5 to +15 ºC

	Temperature compensation
	Integrated NTC network

	Temperature compensation error
	0 °C to 50 °C, 5% relative error

	Nominal sensor life
	> 500,000 % oxygen hours

	Long term output drift
	< 1% volume oxygen per month, typically < -15% over sensor lifetime.

	Shelf life
	24 months

	Load required
	> 10 K Ohms

	Input - membrane protection grille
	No

	Input - hydrophobic membrane
	Yes

	Conformal coating on PCB
	Yes

	Pressure equalization - PCB
	Yes 

	Pressure equalization - output cap
	None

	Weight
	25 g

	Flow diverter
	No

	Packaging
	Gas barrier bag

	Warranty period
	12 months from date of sales invoice

	All specifications are applicable at standard conditions: 1013 mbar, 25°C dry ambient air.



Page 2 of 2

	Drawings/images


[image: image3.png]E Juergensen Marine

o %M%
Distributed by: Email: info@silentthrill.com
Phone: +44 (0) 208 133 6066

Silent Thrill Limited

HH22 Hammerhead Oxygen Sensor

(M| 2013-03 (13 Cé€
SIN JM1123456 ey
® .-

This oxygen sensor contains caustic liquid, in case of leakage avoid contact with eyes
and skin. Do not puncture or staple this bag. Do not remove the oxygen sensor from
this bag until ready to use. Depending upon the ambient conditions the oxygen sensor
may require 30 minutes to reach signal stability after installation.





[image: image4.jpg]15

35

<
A
<

I A

L

M16x1 (5/8-24 UNEF)*

028
29





[image: image5.jpg]Locking Latch

3PIN Molex Connector




[image: image6.jpg]



[image: image7.jpg]



Page 3 of 4

	Labelling


Oxygen sensor label 90.0mm x 15.0mm

White vinyl label, text & graphics – black

[image: image8.jpg]



Gas barrier bag label 63.5mm x 46.6mm

White label, text & graphics – black

[image: image9.png]Vandagraph
>ensor Technologies





NOTE: 6 digit serial number prefixed with: JM1
Page 4 of 4
Vandagraph Sensor Technologies Ltd.

15 Station Road, Cross Hills, Keighley, West Yorkshire BD20 7DT, United Kingdom

Tel +44 (0) 1535 634900   Fax +44 (0) 1535 635582   e-mail: info@vandagraphst.com

www.vandagraphst.com

