[image: image3.jpg]




Specification Datasheet - 8010006

rEvo bvba

CR22D Oxygen Sensor

[image: image4.jpg]@145
|
T
~ i
|
|
| w0
| 3 -
T
|
\
X
! M16 x 1 (5/8-24 UNEF)
@28
1 _—Locking Laich
(-)f%/ TN
h==a
$29

%\N MOLEX CONNECTOR




v1.4
6th July 2016
	Specifications


	VST part number
	8010006

	OEM model number
	CR22D Oxygen Sensor

	Output
	9 – 13 mV

	Output connector
	3 pin Molex

	Input
	Threaded M16 x 1mm

	Input O ring
	Yes

	Range
	0 – 100% Oxygen

	Pressure range
	600 to 2500 mbar

	Repeatability
	<1% volume oxygen at constant temperature and pressure

	Linearity error
	0.1 to 1.0 bar < 3% relative

0.1 to 2.5 bar < 5% relative

(at ppO2 > 1.0 bar after 1 hour equilibration time)

	Warm up time
	< 30 minutes after sensor installed 

	Zero offset voltage
	< 0.5% oxygen reading in 100% nitrogen @ 25ºC after 36 seconds

	Cross interference
	<0.5% volume oxygen response to 5% CO2 balance N2

	Response time
	< 5 s for 90% of final value.

	Operating humidity
	0 – 99% R.H. (non-condensing).

	Influence of humidity
	- 0.03% relative per % R.H. at 25 ºC.

	Influence of mechanical shock
	< 1% after fall from 1m

	Operating temperature range
	0 – 50 ºC

	Storage temperature
	-20 to +50 ºC

	Recommended storage temperature
	+5 to +15 ºC

	Temperature compensation
	Integrated NTC network

	Temperature compensation error
	0 °C to 50 °C, 5% relative error

	Nominal sensor life
	> 500,000 % oxygen hours

	Long term output drift
	< 1% volume oxygen per month, typically < -15% over sensor lifetime.

	Shelf life
	24 months

	Load required
	> 10 K Ohms

	Input - membrane protection grille
	No

	Input - hydrophobic membrane
	Yes

	Conformal coating on PCB
	Yes

	Pressure equalization - PCB
	Yes 

	Pressure equalization - output cap
	None

	Weight
	25 g

	Flow diverter
	No

	Packaging
	Gas barrier bag - clear

	Warranty period
	12 months from date of sales invoice

	All specifications are applicable at standard conditions: 1013 mbar, 25°C dry ambient air.



Page 2 of 4

	Images


[image: image5.jpg]



[image: image6.jpg]



[image: image7.jpg]



[image: image8.png]Vandagraph
>ensor Technologies





Page 3 of 4
	Labelling


Oxygen sensor label 90.0mm x 15.0mm

White vinyl label, text & graphics – black

[image: image1.png]rEvo bvba INSTALL SENSOR PRIOR TO:

CR22D - rEvo OXYGEN SENSOR February 2012

S/N 10612345 Follow manufactrurers
recommendations for

Web: wwnw.info@revo-rebreathers.com
WARNING: CONTAINS CORROSIVE MATERIAL sensor replacement




 
Gas barrier bag label 63.5mm x 46.6mm

White label, text & graphics – black

[image: image2.png]rEvo bvba

CR22D - rEvo OXYGEN SENSOR

SERIAL NO.: 10612345

WARNING!
Sensor contains corrosive electrolyte.
Do not puncture or remove sealed membranes.
Use eye and hand protection if sensor leaks.
Flush and wipe electrolyte from all surfaces.
. Electrolyte can be harmful or fatal if swallowed.
Return sensors damaged in transit to factory.
Do not push on either end of sensor.
Store at reduced temperature to extend life.

f INSTALLATION AND REPLACEMENT OF OXYGEN SENSORS:

ENONPWNS

PLEASE REFER TO YOUR MANUFACTURERS MANUAL FOR
RECOMMENDATIONS AND INSTALLATION INSTRUCTIONS.
INSTALL SENSOR PRIOR TO:
February 2012





NOTE: Serial number format:

1st digit represents the year e.g. 1 = 2011
2nd & 3rd digits represents the month e.g. 06 = June

4th – 8th digits = sequential serial number

Install Sensor Prior To date = Manufactured date + 8 months

Page 4 of 4
Vandagraph Sensor Technologies Ltd.

15 Station Road, Cross Hills, Keighley, West Yorkshire BD20 7DT, United Kingdom

Tel +44 (0) 1535 634900   Fax +44 (0) 1535 635582   e-mail: info@vandagraphst.com

www.vandagraphst.com

