[image: image3.jpg]




Specification Datasheet - 8010015

Uncharted Diving Technology

UDT-22 Oxygen Sensor
[image: image4.jpg]@145
|
T
~ i
|
|
| w0
| 3 -
T
|
\
X
! M16 x 1 (5/8-24 UNEF)
@28
1 _—Locking Laich
(-)f%/ TN
h==a
$29

%\N MOLEX CONNECTOR




v1.1
6th July 2016
	Specifications


	VST part number
	8010015

	OEM model number
	UDT-22 Oxygen Sensor

	Output
	9 – 13 mV

	Output connector
	3 pin Molex

	Input
	Threaded M16 x 1mm

	Input O ring
	Yes

	Range
	0 – 100% oxygen

	Pressure range
	600 to 2500 mbar

	Repeatability
	<1% volume oxygen at constant temperature and pressure

	Linearity error
	0.1 to 1.0 bar < 3% relative

0.1 to 2.5 bar < 5% relative

(at ppO2 > 1.0 bar after 1 hour equilibration time)

	Warm up time
	< 30 minutes after sensor installed 

	Zero offset voltage
	< 0.5% oxygen reading in 100% nitrogen @ 25 ºC after 36 seconds

	Cross interference
	< 0.5% volume oxygen response to 5% CO2 balance N2

	Response time
	< 5 s for 90% of final value.

	Operating humidity
	0 – 99% R.H. (non-condensing).

	Influence of humidity
	- 0.03% relative per % R.H. at 25 ºC.

	Influence of mechanical shock
	< 1% after fall from 1m

	Operating temperature range
	0 – 50 ºC

	Storage temperature
	-20 to +50 ºC

	Recommended storage temperature
	+5 to +15 ºC

	Temperature compensation
	Integrated NTC network

	Temperature compensation error
	0 °C to 50 °C, 5% relative error

	Nominal sensor life
	≥ 500,000 % oxygen hours

	Long term output drift
	< 1% volume oxygen per month, typically < -15% over sensor lifetime.

	Shelf life
	24 months

	Load required
	≥ 10 K Ohms

	Input - membrane protection grille
	No

	Input - hydrophobic membrane
	Yes

	Conformal coating on PCB
	Yes

	Pressure equalization - PCB
	Yes 

	Pressure equalization - output cap
	None

	Weight
	25 g

	Flow diverter
	No

	Packaging
	Gas barrier bag - transparent

	Warranty period
	12 months from date of sales invoice

	All specifications are applicable at standard conditions: 1013 mbar, 25 °C dry ambient air.



Page 2 of 4

	Drawings/images


[image: image5.jpg]



[image: image6.jpg]



[image: image7.png]


[image: image8.png]Vandagraph
>ensor Technologies





Page 3 of 4

	Labelling


Oxygen sensor label: 90.0mm x 15.0mm

White vinyl label, text & graphics – black

[image: image1.jpg]www.divetechnology.com UDT-22 Oxygen Sensor

w UNCHARTED
Y U DT | lTlgllrlll;uwﬁv A @ TXXXXXX
g CE o HIII\IIII\IIIIIIHI\||||||||||||\II\IHI\IHI\IHIII




Gas barrier bag label: 63.5mm x 46.6mm

White label, text & graphics – black

[image: image2.jpg]N/

- UHENARIED Email: octagon@unchartereddiving.com
Tilkougey wwwdivetechnology.com

UDT-22 Oxygen Sensor

&I 2015-11
SIN  UDTxxxxxx

T —ee A

This oxygen sensor contains caustic liquid, in case of leakage avoid contact with eyes
and skin. Do not puncture or staple this bag. Do not remove the oxygen sensor from
this bag until ready to use. Depending upon the ambient conditions the oxygen sensor
may require 30 minutes to reach signal stability after installation.





Serial number prefix: UDT 

Page 4 of 4
Vandagraph Sensor Technologies Ltd.

15 Station Road, Cross Hills, Keighley, West Yorkshire BD20 7DT, United Kingdom

Tel +44 (0) 1535 634900   Fax +44 (0) 1535 635582   e-mail: info@vandagraphst.com

www.vandagraphst.com

