[image: image10.png]


 

                                     Nufer Radiant Warmer

COMPANY OPERATING PROCEDURES

Nufer Radiant Warmer Test Procedure

(0310040) – Mobile Stand 

VM3/COP/53.01

Date:
  24-Jan-03
Revision date: 10-Jan-08
Issue:   3





[image: image1.jpg]



Equipment Required:
Rigel 255 safety tester, stapler, tape gun & tape, knife, philips screwdriver, flat blade screwdriver, M4 allen key,

[image: image2.png]


 
Warning: 

Rigel 255 safety tester: A hazard to the user is present when the safety tester is carrying out its automatic test sequence through exposure to the mains voltage supply.  The operator should be suitably trained before using the equipment.

· Ensure Correct Plug Fitted for Intended Market

· Only Skilled / Competent personnel are allowed to use the Rigel Safety Tester.

· Do not touch or allow any other person to touch the equipment whilst the safety tester is testing.  

· Do not leave the safety tester unsupervised whilst carrying out its’ automatic test sequence.  

· Do not touch the radiant warmer element, reflector or guard bars when hot.

[image: image10.png]Power Level Indicators

Light, on /off Switch
Radiator 

on / off Switch

Select Lower Heat Level

[image: image3.wmf]|

0

CERATHERM 600

2

1

3

4






Select Higher Heat Level

Heating Indicator

Alarm Indicator & Cancel




Fuse Holder (2 x 3.15 AT)




[image: image4.wmf]





Input Socket for Mains Cable (220-240V AC, 50Hz)


Functional & Safety Test Method:

1. Ensure that all Nufer logos, tape and markings are removed from the packaging.

2. Unpack the box and ensure that all parts have been delivered in accordance with the contents checklist (VM/COP/53.02).  Ensure all parts are free from physical defects.  Remove any marks as necessary. Complete the checklist and photocopy.

3. Remove and discard the “Nufer Medical” label from the front of the radiant warmer and apply a Viamed label (110mm x 22mm – pt no 0330130) centrally above the control panel.


[image: image5.jpg]




[image: image6.jpg]VIAMED





4. Place the radiant warmer element side down onto the workbench.

5. Attach the radiant warmer  stand adaptor to the radiant warmer. Tighten each of the four bolts using the M4 Allen key. Connect and bolt the adaptor to the mobile stand unit and tighten.


[image: image7.jpg]





[image: image8.jpg]b





6. Connect the radiant warmer to the mains outlet, using the mains lead.




[image: image9.jpg]IVIAMED

| %

—





7. Switch the radiant warmer on at the control panel and check that the halogen light works when switched on at the control panel.  Switch the halogen light off.

8. Press “ (<) ” & “ (>) ” to cycle up and down the heat levels.

9. The heating indicator shows when the radiant warmer element is energized (indicator lit) to when it is not (indicator unlit). Ensure that the length of time the heating indicator is lit to when it is not reflects the heat level setting. 

10. Select heat level .

11. Monitor the radiant warmer and ensure the unit goes into alarm after 15 minutes.

12. When in alarm, ensure that the radiant warmer heat level drops to the safety setting of level , the alarm sound and the alarm indicator / cancel flashes.

13. When in alarm, press the alarm indicator / cancel.  The heat level should return to level , the alarm indicator should stop flashing and the alarm should stop sounding.

14. With the radiant warmer still switched on and hot, disconnect the mains plug from the mains outlet and connect to the 3-pin socket on the Rigel 255 safety tester.

15. Connect the earth return lead to one of the radiant warmer grab rail handles. 

16. Ensure the heater passes the safety test.

17. Print & bind current versions of the Radiant Warmer Operators Manual (PN 0391000) and Servicing Manual (PN 0391010).

18. Repackage all parts and place the manuals, one copy of the safety test certificate and one copy of the contents sheet into the box.

19. Secure the box adequately for transport.  Lightly tape the second contents sheet to the box lid.

20. Ensure stock sheets are printed, completed & filed.

21. Staple the second Viamed safety certificate to the Nufer safety test certificate and file.

Data Entry into Rigel 255 Safety Tester:

1. System Test Results: Results correct? : Press “Y”.

2. Enter Name or Initials: Press multifunction key 1: “Enter”.

3. Press multifunction key 2: “DATA”.  Press multifunction key 5: “EQUIP”.

4. Type the equipment name: “HEATER XXXXXXXX” where XXXXXXXX is the radiant warmer serial number

5. Unknown equipment No. Add to data stores (Y/N)? : Press “Y”.

6. Equipment configuration record: 


Type the following entries, following each with keyboard Enter:


Location: “VIAMED”


Ward: “QA”


Manufacturer: No entry


Description: “INFANT RADIANT WARMER”, 


Serial No.: “XXXXXXXX” where XXXXXXXX is the radiant warmer serial number


Class: “1”


Type: “B”


Applied parts: “0”


Test Interval: “12”


Test mode: “A”


Test name: “XX…XX” where XX…XX is the testers name or initials


Press multifunction key 1: “ENTER”.

7. Save data: Press “Y”.

8. Press “◄” twice to get to “TEST SELECTION MENU”

9. Press multifunction key 5: “AUTO”.

10. In AUTOMATIC TEST MODE, type in the equipment name, “HEATER XXXXXXXX” where XXXXXXXX is the radiant warmer serial number

11. Press multifunction key 1: “ENTER”.

12. Press start on the safety tester (|). (To stop the safety test at any point press .  The automatic test sequence will stop after it has completed its current step). 

13. After completion of the automatic test sequence, in the Additional Notes box, type, “Safety test carried out with a mobile stand”.

14. Press multifunction key 1: “ENTER”.  Press multifunction key 2: “REMOTE”.

15. Ensure the calibration certificate (QC10) is printed and photocopy.

16. Disconnect the safety tester and store the earth return lead.

Page 6 of 6

