Technical Data Sheet   Oxygen Sensor R-17


[image: image5.png]IMVIAMED




[image: image2.jpg]



[image: image3.jpg]




Specifications

Output – With 10K Load
10  3mV
Cross Interference

Response time :step change 90%
<6 seconds
Less than 1.25% O2 

Response to:

In a mixture of:

30% O2 / 70%N20
6.0% Halothane

Accuracy
 1.0% O2

6.0% Isoflurane

Electrical Interface
Connectors

6.0% Enflurane

Expected life *
36 Months

7.0% Sevoflurane

Storage Temperature
0 - 50C

20.0% Desflurane

Operating Temperature 
0 - 40C



Operating Humidity
Up to 99% RH
Warranty**
12 months

Cross Reference Chart

OEM 
Reference Number
Equipment

Analytical Industries
PSR-11-917-J


Biotest
406931, C102-1


BPR
ST-11


Catalyst Research
406931
Mini-Ox I, II & III

Ceramatec
CAG-13


City-tech
MOX-2, MOX-02


Dameca
42035-46


Envitec
OOM102-1


Gotek
PSR-11-917-J


Hamilton
51250
Raphael

IT Gambert
M-03


MSA
406931, 468324
Mini-Ox I, II & III, IA & Mini-Ox 3000

Penlon
58780
Anaesthetic Machines

Puritan Bennet
4-020933-00
7200

Sensortec
ST-11


Spacelabs
015-0285-00


VTI
100580


Viamed has been supplying oxygen sensors specifically designed for medical use for 28 years and continues to offer the widest range of state of the art sensors. 

For further information either technical or application please do not hesitate to contact us.

Almost all information can be supplied electronically if required.

*Expected life in Air at sea level Standard temperature and pressure calculated assuming a maximum 70% anode usage. For 100% anode use add 30% to life figures

** Non-quibble Viamed warranty 12 months, Limited warranty due to manufacturing default 2 years.

[image: image1.jpg]


[image: image4.png]Viamed Limited. 15 Station Road, Cross Hills,
Keighley, West Yorkshire BD20 7DT, United Kingdom
Tel: +44 (0)1535 634542  Fax: +44 (0)1535 635582

E-mail: info@viamed.co.uk Web site: www.viamed.co.uk





_176921852.bin

