[image: image1.jpg]3689402

= AEEE1R0.




VANDAGRAPH

TEK-OX T

TEK-OX Assembly Procedure

VM3/COP/42.01

Date: 02/02/08
Revision Date: 4-Feb-08
Issue: 1


Parts Required

Quantity
Description
Part No.


1
TEK-OX Black  Front Cover
9530005


1
TEK-OX Black  Bottom cover
9530004


1
Blu-Ox PC Lens

9530009


1
TEK-OX Black Shell left
9530012


1
TEK-OX Black  Shell Right
9530013


1
Blu-Ox Base Seal
9530002


1
Blu-Ox Front gasket
9530006


1
Blu-Ox  switch
9730036


1
Blu-Ox Cal contro
9730040


1
Blu-Ox Photo thread
9530010


1
Blu-Ox key ring
9530008


1
Blu-Ox Inside label
9530007


1
Blu-Ox Instruction manual
9590001


1
Blu-Ox  switch waterproof kit
9730037


1
Blu-Ox battery terminals right



1
Blu-Ox battery terminals left



1
LCD assembly



1
PCB assembly



1
Cable assembly



2
Blu-Ox batteries AA
9060005


1
R-17VAN
9711017



Screws Black




Screws Black








`1
Locktite O2 compatible



1
Silicon grease



1
Slilicone sealant Black
9730004


[image: image2.jpg]Green 90mm Jack centre \ _
Brown 90mm Jackgnd Black 80mm Switch
White 80 mm Switch

Orange 80mm Pot cntr
Yellow 80mm Pot Left

Violet 80mm Potright

Red 90mm Batt +ve
Blue 90mmBatt -ve

Cable sub assembly PA 29/11/2007




[image: image3.jpg]Long cntr Orange





[image: image4.jpg]T AA

]E:é MN1500





Assembly:

1. Cut, strip & tin cable assembly wires

2. [image: image5.jpg]Sensor R-17 VAN
Batt: MN1500 x2

Ser. No. TX




Feed Black & white through the left hand side of enclosure and ensure the O ring is in place on the switch body. Solder wires to switch

3. Lubricate with silicone grease the O ring

4. Feed the switch into the enclosure mate the locating peg and add the bridge sealing kit,

5. This has to be tightened through the right hand enclosure aperture

6. Add nut & locking washer washer  to Orange Yellow & Violet wires

7. Feed wire through main body Right hand hole

8. Add O ring

9. Solder wires to the pot see Diagram

10. Feed pot into enclosure body add Permabond
MH02 and tighten nut.

11. Feed Red & Blue wires through Battery terminal access holes
Red Right  Blue Left and solder to correct terminal. They only work one way

12. Feed the terminals up into the enclosure.. 

13. Cut back the terminals of the Jack plug to about 1cm.

14. Feed Jack plug into enclosure  using nut to secure. Use Permabond

15. Solder Green to centre and Brown to outside.

16. Add the connector to the PCB.

17. Connect the LCD to the PCB

18. Stick the LCD to the PCB with double sided tape.

19. Ensure alignment is good

20. Add Batteries and fix the base on with one screw

21. Test the unit

22. Remove the batteries

23. Add Black silicone rubber inside enclosure to 
Jack Plug
Switch
Calibration Knob

24. Locate LCD & PCB into enclosure. Sockets at the bottom

25. Add a thin layer of Silicone grease to both sides of the front panel seal

26. Locate the seal on the back of the front panel

27. [image: image6.jpg]TEK-OX Oxygen Analyser
VANDAGRAPH LTD. d
15 Station Read, Gross Hills

Keighlay, W.Yarkshirz, BD20 70T, UK
Tel. +44 (0)1535 624300

Fax. +44 {0)1535 6355582 c €
Sen Na. TX _—




Offer the front panel up to the enclosure

28. It locates with small fragile lugs on each side and a lock on the top.

29. Screw the front panel on with 2 size ?? screws

30. [image: image7.jpg]


Add the two outer covers 

31. Inside add the Battery labels

32. +ve Right

33. - ve Left

34. Add serial number label  white to inside 

35. [image: image8.jpg]


Add same serial number but Black to rear of Instrument

36. Lubricate the battery compartment seal and insert into the battery compartment cover

37. Add Sensor

38. Add batteries and fix cover with three ?? screws

39. Check the unit functions

40. Add the front labels 

41. Switch On

42. Check correct software version

43. Check correct title

44. Check the correct colour switch 

45. Yellow auto switch off Green no auto switch off

46. Set Cal 20.9%

47. Check 100%

48. Add Foam to box

49. Add label to box

50. Add Quick-OX or DINKIT

51. Add user guides
TEK-OX
Quick-OX

Equipment Required:


Soldering Iron: side cutters: Spanner;Long nose pliers











VM3COP42.01 TEK-OX Assembly 080201.doc         July 17, 2006      Page 4 of 4

