[image: image14.png]


 

                      SpO2 Extension Cable Assembly

	COMPANY OPERATING PROCEDURES


0019248
    SPACELABS   
P924E8

VM3/COP/33.00
	Date:
  5-Nov-02
	Revision date: 14-Feb-05
	Issue:   3




[image: image1.wmf]D-Type Male 10 Pin

2.4m

Ø

6 x 43 Clear 

heatshrink. For serial 

no.

D-Type Female 9 Pin


[image: image2.jpg]


                                                                                    [image: image3.jpg]



Equipment required:
Soldering iron (0060120), solder (0050012), Wire stripper (0060030), Flush Cutter (0060010), Snipe nose pliers (0060021), ‘helping hand’ (0060145), Heat gun (0060100).

Parts list:
 

	D-Type male 9-pin side
	D-Type female 9-pin side

	Qty
	Description
	Part No.
	Qty
	Description
	Part No.

	
	
	
	
	
	

	1
	10 pin socket kit (spacelabs)
	0010699
	1
	P956E8 Extension Cable
	0019568

	(1)
	
[image: image4.wmf]     Pin housing
	kit
	
	
	

	(1)
	
[image: image5.wmf]   Upper casing
	kit
	
	
	

	(1)
	
[image: image6.wmf]    Lower Casing
	kit
	
	
	

	(1)
	
[image: image7.wmf] Strain relief
	kit
	
	
	

	(1)
	       
[image: image8.wmf]              Cable grip
	kit
	
	
	

	1
	(6 x 43mm Clear heat shrink
	0032331

(roll)
	
	
	

	1
	(1.6 x 11mm heat shrink
	0032310
(roll)
	
	
	

	1
	     
[image: image9.wmf]R1

          60K4 ( Resistor
	0032130
	
	
	

	1
	(6 x 10mm heat shrink
	0032321
(roll)
	
	
	

	
	
	
	
	
	

	
	
	
	
	
	


ASSEMBLY OPERATIONS

1.
Take a 0019568 pre-manufactured cable, which requires the 9 pin male connector to be cut off and replaced with a 10-pin Spacelabs connector.

D-Type male 9-pin side:


[image: image10.wmf]W

(3) 

Combined 

outer and 

inner 

shields 

(is+ os) 


[image: image14.png]Fig 2.2  
1.
Feed (6 x 10mm heat shrink (clear), strain relief and cable grip over the end of the cable.

2.
Strip 30mm off outer jacket of wire to reveal coloured wires, outer shield, and nylon/paper wire packing.

3.
Cut all packing to the base and the resistor legs 8mm and 18mm.

4.
Strip 30mm off inner jacket to reveal black and white wires and the inner shield.

5.
Twist outer and inner shields together.

6.
Trim orange and shields to 15mm.

7.
Strip jacket of every wire 2mm to reveal copper core.

8.
Heat (1.6 x 11mm heat shrink over the twisted shields and orange wireto cover naked wire.

9.
Apply small amount of solder to ends of each wire and shields.

10.
Solder the orange wire together with the combined shields to the resistor leg on the side of pin location 3.

11.
Solder ends of every wire and shields to the correct positions on the rear of the connector.

12.
Clamp cable grip approximately 2mm from end of outer jacket.

13.
Place (6 x 10mm heat shrink over cable grip and beginning of wires and heat to shrink firmly over.

14.
Fill the rear of the connector with hot glue, then push the outer casing over cable grip and wires to fit around the pin housing.

TESTING 

1.
Plug an 856RA finger probe into the extension cable, and use a multimeter to check between the pins to achieve the following results:

	Pin 6 to Pin 1
=
0.5v

Pin 4 to Pin 5
=
1.1v (I.R)

Pin 5 to Pin 4
=
1.5v (Red)

Pin 2 to Pin 3
=
60.4 (k()

Pin 3 to Pin 7
=
7.5 (k()
	
[image: image11.wmf]5

4

3

2

1

6

7

8

9

10




2.
Check the cable is of correct quality standard. (See VM/COP/30.11 for details).

Labelling

1.
Labels:
 to be attached facing upwards as looking at the top of the probe.


· 1 x CE Label.

· 1 x Serial no. Label.

· 1x Orange ‘Do Not Throw Away’ Label (correct one of two is dependant of country unit is being sold to).

[image: image12.wmf]Do Not

Throw

Away

Orange dumbell 

labell

CE

0086

Ø6 x 43 clear heat shrink

Serial No.


Quality Assurance (QA)

1.
Plug an 856RA finger probe into the extension cable, and use a multimeter to check between the pins to achieve the following results:

	Pin 6 to Pin 1
=
0.5v

Pin 4 to Pin 5
=
1.1v (I.R)

Pin 5 to Pin 4
=
1.5v (Red)

Pin 2 to Pin 3
=
60.4 (k()

Pin 3 to Pin 7
=
7.5 (k()
	
[image: image13.wmf]5

4

3

2

1

6

7

8

9

10




2.
Check the cable is of correct quality standard. (See VM/COP/30.11 for details).

3.
Fill and sign attached paperwork.

Packaging 

1.
Visually check all labels are attached properly

2.
Using a twist tie (bunny clip) wrap the cable and place in a small blue Viamed plastic box, ensuring the cable is inserted in a neat and tidy presentable manor.

3.
Place a serial number sticker (supplied with the batch) on the front face of the box.

4.
Place a packed and tested sticker (also containing initials of the individual who is packing) on the right hand side top left corner of the box. Do not close box.

Final QA

1.
Final inspection. Visually ensure cable sit neatly within the box and is in a presentable state.

2.
Boxes are ready to stock in stores.




(2,3) R1 = 60k4














Page 2 of 4

_1098007563.dwg

_1167831610.dwg

_1105962007.dwg

_1105962046.dwg

_1098009614.dwg

_1098007421.dwg

_1098007503.dwg

_1070369566.dwg

_1098007235.dwg

_1069494359.dwg

