[image: image27.png]


 

                      SpO2 Extension Cable Assembly

	COMPANY OPERATING PROCEDURES


0019252
SPACELABS
P925E2

VM3/COP/33.01
	Date:
  11-Dec-01
	Revision date: 19-Sep-03
	Issue:   3




[image: image1.wmf]Hypertronics Female 7 Pin

Hypertronics Male 12 Pin

0.6M

Enlarged front view

Enlarged front view


 [image: image2.jpg]


                                                                [image: image3.jpg]



Equipment required:
Soldering iron (0060120), solder (0050012), Wire stripper (0060030), Flush Cutter (0060010), Snipe nose pliers (0060021), ‘helping hand’ (0060145), Heat gun (0060100).

Parts list:
Kit and parts required.

	Hypertronics Female

7-pin Side
	Hypertronics male 

12-pin Side

	Qty
	Description
	Part No.
	Qty
	Description
	Part No.

	
	
	
	
	
	

	1
	Hypertronics Female 

7-pin Kit
	0010606
	1
	Hypertronics male 

12-pin Kit
	0010602

	(1)
	
[image: image4.wmf]         Pin Housing
	kit
	(1)
	  
[image: image5.wmf]         Pin Housing
	kit

	(9)
	
[image: image6.wmf]          Pins
	kit
	(12)
	
[image: image7.wmf]           Pins
	kit

	(1)
	       
[image: image8.wmf]               Cable grip
	kit
	(1)
	       
[image: image9.wmf]              Cable grip
	kit

	(1)
	  
[image: image10.wmf]       Strain relief
	kit
	(1)
	  
[image: image11.wmf]       Strain relief
	kit

	(1)
	
[image: image12.wmf]            Colett
	kit
	(1)
	   
[image: image13.wmf]            Collett
	kit

	(1)
	
[image: image14.wmf]           Barrell
	kit
	(1)
	  
[image: image15.wmf]           Barrel
	kit

	1
	0.7m of standard 6-core cable
	0030513
(roll)
	1
	
[image: image16.wmf]R1

     47.5 k( Resistor
	0032110

	1
	(1.6 x 17mm heat shrink
	0032310

(roll)
	1
	(6 x 43mm Clear heat shrink
	0032331

(roll)

	1
	(6 x 10mm heat shrink
	0032321

(roll)
	1
	(6 x 10mm heat shrink
	0032321

(roll)

	
	
	
	
	
	


ASSEMBLY OPERATIONS

1.
Pre Heat soldering iron temperature to 240oc.

2.
Collect all required parts and equipment listed above.

3.
Cut a 0.7 metre length of standard 6-core cable. (Details shown below).

                           Fig 1.  
[image: image17.wmf]1.2 Cable Cross Section

1.0 Mechanical Specifications

1.1 Cable Schematic

Black

White

Blue

Inner Shield

Orange

Yellow

Red

Outer Shield

Outer Jacket

Outer Shield

Red,Orange, Yellow, Blue

Inner Shield

Black, White Twisted Pair

Twisted Pair

Paper/ nylon packing to make round


Hypertronics female 7 pin side:

Fig 2.1
[image: image18.wmf]Pin Housing

Cable grip

Strain relief

Collett

Barrell

Pin Housing Rear View

2

1

6

5

7

3

4

Pins

(2) Red

(-) Blue

(6) Yellow

(1) Orange

(3) Black

(4) White

(5) 

Combined 

outer and 

inner 

shields 

(is+ os) 


1.
Feed barrel, collett, strain relief, cable grip and (6 x 10mm heat shrink onto the cable.

2.
Strip 20mm off outer jacket of wire to reveal 4 coloured wires, outer shield, and nylon/paper wire packing.

3.
Cut all packing, and blue wire to the base.

4.
Strip 20mm off inner jacket of wire to reveal black and white wires and the inner shield.

5.
Twist outer and inner shields together.

6.
Trim ends of wires and shields to the same length.

7.
Strip jacket of every wire 2mm to reveal copper core.

8.
Apply small amount of solder to ends of each wire and shields, then solder each of the 5 wires to the rear separate pins.

9.
Place (1.6 x 17mm heat shrink over twisted inner and outer shields to cover naked wire, and solder to the rear of a pin.

10.
Clamp cable grip approximately 2mm from end of outer jacket.

11.
Place (6 x 10mm heat shrink over cable grip and beginning of wires and heat to shrink firmly over.

12.
Insert pins into correct locations (as shown in fig 2.1) and push firmly into place.

13.
Push strain relief up to cable grip, collett over strain relief up to the pin housing (locate in keyway), hold the pin housing fixed and screw the barrel to the pin housing.

Hypertronics male 12 pin side:

Fig 2.2  
[image: image19.wmf](3) 

Combined 

outer and 

inner 

shields 

(is+ os) 

(9) Red

Pins

End Cap

Collett

Strain Relief

Pin Housing

Barrell

Cable grip

1

2

3

4

5

6

7

9

8

10

11

12

Pin Housing Rear View

R1

W

k

(7,3) R1 = 47.5

(2) Black

(1) White

(10) Orange

(12) Yellow

(-) Blue


1.
Feed (6 x 43mm (clear) heat shrink, end cap, collett, strain relief, (6 x 10mm (black) heat shrink and cable grip onto the cable.

2.
Strip 20mm off outer jacket of wire- to reveal 4 coloured wires, outer shield, and nylon/paper wire packing.

3.
Cut all packing, and blue wire to the base.

4.
Strip 20mm off inner jacket of wire- to reveal black and white wires and the inner shield.

5.
Twist outer and inner shields together and trim to approximately 8mm long.

6.
Strip jacket of every wire 2mm to reveal copper core.

7.
Apply small amount of solder to ends of each wire and shields.

8.
Trim one of the legs of the resistor to 4mm and the other to 15mm .

9.
Solder each leg of the resistor into the rear of 2 separate pins and push pins firmly into correct locations as shown in   fig 2.2.

10.
Solder all 5 wires to the rear of 5 pins and push/pull firmly into correct locations shown in     fig 2.2.

11.
Solder shields to the leg of the resistor on the pin location 3 side.

12.
Clamp cable grip approximately 2mm from outer jacket end.

13.
Place (6 x 10mm heat shrink over cable grip and beginning of wires and heat to shrink firmly around.

14.
Push the strain relief up to the cable grip, collett over the strain relief up to the pin housing, and into the barrel and finally screw end cap onto the barrel.

TESTING

1.
Attach male Hypertronics 12-pin side to an 877 Spacelabs adapter cable and then to the test box connector labelled (B).

2.
Attach female Hypertronics 7-pin side to a Novametrics finger probe.

3.
Check display is showing correct characteristic as shown below. (At correct switch positions)

[image: image20.jpg]



[image: image21.jpg]



[image: image22.jpg]




     Pos 2. LED


Pos 3. IR

         Pos 4. Detector

4.
‘Play’ with wire at connections to see if any change in the display (i.e. flickering etc).

5.
If there is any movement of signal, the extension wire must be taken apart and all connections checked and re-soldered. Then tested again, until results are satisfactory.

6.
Check the cable is of correct quality standard. (See VM/COP/30.11 for details).

7.
Connect male 12-pin side to a test cable then to the Novametrix monitor, and attach probe to the finger. Check SpO2 level is around 95-100.

Labelling

1.
Labels:
to be attached facing upwards as looking at the top of the probe.


· 1 x CE Label

· 1 x serial no. Label

· 1x Orange ‘Do Not Throw Away’ Label (correct one of two is dependant of country unit is being sold to).

[image: image23.wmf]Serial No.

Ø6 x 43 clear heat shrink

Do Not

Throw

Away

CE

0086

Orange dumbell 

labell


Quality Assurance (QA)

1.
Attach male Hypertronics 12-pin side to an 877 Spacelabs adapter cable and then to the test box connector labelled (B).

2.
Attach female Hypertronics 7-pin side to a Novametrics finger probe.

3.
Check display is showing correct characteristic as shown below. (At correct switch positions)

[image: image24.jpg]



[image: image25.jpg]



[image: image26.jpg]




     Pos 2. LED


Pos 3. IR

         Pos 4. Detector

4.
‘Play’ with wire at connections to see if any change in the display (i.e. flickering etc).

5.
If there is any movement of signal, the extension wire must be taken apart and all connections checked and re-soldered. Then tested again, until results are satisfactory.

6.
Check the cable is of correct quality standard. (See VM/COP/30.11 for details).

7.
Connect male 12-pin side to a test cable then to the Novametrix monitor, and attach probe to the finger. Check SpO2 level is around 95-100.

8.
Fill and sign attached paperwork.

Packaging 

1.
Visually check all labels are attached properly

2.
Using a twist tie (bunny clip) wrap the cable and place in a small blue Viamed plastic box, ensuring the cable is inserted in a neat and tidy presentable manor.

3.
Place a serial number sticker (supplied with the batch) on the front face of the box.

4.
Place a packed and tested sticker (also containing initials of the individual who is packing) on the right hand side top left corner of the box. Do not close box.

Final QA

1.
Final inspection. Visually ensure cable sit neatly within the box and is in a presentable state.

2.
Boxes are ready to stock in stores.
















Page 6 of 6

[image: image27.png]_1069495046.dwg

_1069844807.dwg

_1069844975.dwg

_1071045612.dwg

_1091425918.dwg

_1071045304.dwg

_1069844908.dwg

_1069504476.dwg

_1069744209.dwg

_1069763934.dwg

_1069495127.dwg

_1069494359.dwg

_1069494489.dwg

_1069494795.dwg

_1069494204.dwg

