[image: image16.png]


 

                      SpO2 Extension Cable Assembly

	COMPANY OPERATING PROCEDURES


0019618   
   BCI   
P961E8

VM3/COP/33.09

	Date:
  17-Dec-01
	Revision date: 19-Sep-03
	Issue:   3




[image: image1.wmf]Enlarged front view

Redel Male 7 Pin

Ø

6 x 43 Clear 

heatshrink. For serial 

no.

P956E8 Extension cable

[image: image2.jpg]







[image: image3.jpg]



Equipment required:
Soldering iron (0060120), solder (0050012), Wire stripper (0060030), Flush Cutter (0060010), Snipe nose pliers (0060021), ‘helping hand’ (0060145), Heat gun (0060100).

Parts list:
Kit and parts required. (Continued over page)

	D-Type female 9-pin Side
	Redel male 7-pin Side

	Qty
	Description
	Part No.
	Qty
	Description
	Part No.

	
	
	
	
	
	

	1
	D-type extension cable

(Female side)
	0019568
	1
	Redel male 7-pin kit
	0030795

	
	
	
	(1)
	
[image: image4.wmf]                Sleeve
	Kit

	
	
	
	(1)
	
[image: image5.wmf]             Collett
	Kit

	
	
	
	(1)
	
[image: image6.wmf]         Connector Pin
	Kit

	
	
	
	(1)
	
[image: image7.wmf]      Outer Sleeve
	Kit

	
	
	
	(1)
	
[image: image8.wmf]Strain Relief (orange)
	0030794

	
	
	
	1
	(6 x 43mm Clear heat shrink
	0032331

(roll)

	
	
	
	1
	(6 x 10mm heat shrink
	0032321

(roll)

	
	
	
	1
	(1.6 x 17mm heat shrink
	0032310

(roll)


ASSEMBLY OPERATIONS

1.
Pre Heat soldering iron temperature to 240oc.

2.
Collect all required parts and equipment listed above.

D-type female 9-pin side:


[image: image9.wmf]2.4m

D-Type Female 9 Pin

Ø

6 x 43 Clear 

heatshrink. For serial 

no.

D-Type Male 9 Pin


 Fig 2.1.  0019568 Extension Cable.

1.
The 0019568 cable is already assembled. The female side does not need to be modified, but the male side needs to be cut off ready for the Redel male 7-pin connector to be fitted.


Redel male 7 pin side:

Fig 2.2  
[image: image10.wmf](2) 

Combined 

outer and 

inner 

shields 

(is+ os) 

7

(1) Black

(6) Yellow

(7) Orange

(5) Red

(4) White

(3) Blue

Strain Relief

Sleeve

Collett

Outer Sleeve

Conector Pin

5

6

3

4

2

1

Connector Rear View


1.
Feed (6 x 43mm (clear) heat shrink, strain relief, sleeve, collett, (6 x 10mm (black) heat shrink over end of cable.

2.
Strip 10mm off outer jacket of wire to reveal coloured wires, outer shield, and nylon/paper wire packing.

3.
Cut all packing to the base.

4.
Strip 10mm off inner jacket of wire- to reveal black and white wires and the inner shield.

5.
Twist outer and inner shields together.

6.
Trim (tidy) ends of all wires and shields to the same length.

7.
Strip jacket of every wire 2mm to reveal copper core.

8.
Apply small amount of solder to ends of each wire and shields.

9.
Heat (1.6 x 7mm heat shrink over twisted shield pair to insulate.

10.
Solder wires and paired shields to the rear of the connector pins as shown in fig 2.2.

11.
Push collett up to the connector pin, and screw the sleeve into the outer sleeve.

TESTING 

1.
Attach the Redel side to a test lead and then the test box connector marked ‘A’.

2.
Attach D-type side to a BCI finger probe.

3.
Check display is showing correct characteristics as shown below. (At correct switch positions)

[image: image11.jpg]


                    
[image: image12.jpg]R -

—— TN m—— T~

e

A AR e




                     NB: Gap to be at bottom of display


            
     Position 2. IR, LED.
       
                           Position 4. Detector

4.
If the gap is at the top of the screen then the LED is wired incorrectly.

5.
‘Play’ with wire at connections to see if any change in the display (i.e. flickering etc).

6.
If there is any movement of signal, the extension wire must be taken apart and all connections checked and re-soldered. Then tested again until results are satisfactory.

7.
Check the cable is of correct quality standard. (See VM/COP/30.11 for details).

8.
Connect Redel 7-Pin side to a BCI 71000A2 portable monitor and attach probe on finger to check SpO2 level. (Ideal reading 95-100.)

Labelling

1.
Labels: to be attached facing upwards as looking at the top of the probe.


· 1 x CE Label

· 1 x serial no. Label

· 1x Orange / Brown BCI Label 


[image: image13.wmf]Orange/Brown 

dumbell labell

For use on

BCI

Compatible

Pulse Oximeters

0086

CE

Ø6 x 43 clear heat shrink

Serial No.


Quality Assurance (QA)

1.
Attach the Redel side to a test lead and then the test box connector marked ‘A’.

2.
Attach D-type side to a BCI finger probe.

3.
Check display is showing correct characteristics as shown below. (At correct switch positions)

[image: image14.jpg]


                    
[image: image15.jpg]R -

—— TN m—— T~

e

A AR e




                     NB: Gap to be at bottom of display


            
     Position 2. IR, LED.
       
                           Position 4. Detector

4.
If the gap is at the top of the screen then the LED is wired incorrectly.

5.
‘Play’ with wire at connections to see if any change in the display (i.e. flickering etc).

6.
If there is any movement of signal, the extension wire must be taken apart and all connections checked and re-soldered. Then tested again until results are satisfactory.

7.
Check the cable is of correct quality standard. (See VM/COP/30.11 for details).

8.
Connect Redel 7-Pin side to a BCI 71000A2 portable monitor and attach probe on finger to check SpO2 level. (Ideal reading 95-100.)

9.
Fill and sign attached paperwork.

10.
Test 10 % of batch on DL3000 simulator.

11.
Log all results on compatibility sheet.

Packaging 

1.
Visually check all labels are attached properly

2.
Using a twist tie (bunny clip) wrap the cable and place in a small blue Viamed plastic box, ensuring the cable is inserted in a neat and tidy presentable manor.

3.
Place a serial number sticker (supplied with the batch) on the front face of the box.

4.
Place a packed and tested sticker (also containing initials of the individual who is packing) on the right hand side top left corner of the box. Do not close box.

Final QA

1.
Final inspection. Visually ensure cable sit neatly within the box and is in a presentable state.

2.
Boxes are ready to stock in stores.
















Page 4 of 5

[image: image16.png]_1070101265.dwg

_1070179703.dwg

_1091450676.dwg

_1070454204.dwg

_1070101508.dwg

_1070101104.dwg

_1070101179.dwg

_1070101036.dwg

_1070100912.dwg

