[image: image23.png]


 

                      SpO2 Extension Cable Assembly

	COMPANY OPERATING PROCEDURES


0019630   
     DATASCOPE   
P963E10

VM3/COP/33.10
	Date:
  18-Dec-01
	Revision date: 10-Sep-04
	Issue:   3




[image: image1.wmf]Ø

6 x 43 Clear 

heatshrink. For serial 

no.

3m

Enlarged front view

D-Type Female 9 Pin

DIN Male 8 Pin


[image: image2.jpg]


                                                        [image: image3.jpg]



Equipment required:
Soldering iron (0060120), solder (0050012), Wire stripper (0060030), Flush Cutter (0060010), Snipe nose pliers (0060021), ‘helping hand’ (0060145), Heat gun (0060100).

Parts list:
Kit and parts required. (Continued over page)

	D-type female

9-pin side
	Din male

8-pin side

	Qty
	Description
	Part No.
	Qty
	Description
	Part No.

	
	
	
	
	
	

	1
	D-type female 9-pin kit
	0010760
	1
	Din male 8-pin kit
	0010790

	(1)
	
[image: image4.wmf]  Outer Casing
	kit
	(1)
	
[image: image5.wmf]         Barrel
	kit

	(1)
	
[image: image6.wmf]           Pin Housing
	kit
	(1)
	
[image: image7.wmf]          Cable grip
	kit

	(9)
	
[image: image8.wmf]           Pins
	kit
	(1)
	
[image: image9.wmf]             Connector
	kit

	1
	       
[image: image10.wmf]              Cable grip
	010500

(pack of 100)
	1
	
[image: image11.wmf]R1

     22.1 ( Resistor
	0032020

	1
	3.1m 6-core cable
	0030513

(roll)
	1
	
[image: image12.wmf]R2

  47.5 k( Resistor
	0032110

	1
	(1.6 x 17mm heat shrink
	0032310

(roll)
	1
	(6 x 43mm Clear heat shrink
	0032331

(roll)

	1
	(6 x 10mm heat shrink
	0032321

(roll)
	1
	(6 x 40mm heat shrink
	0032321

(roll)

	
	
	
	1
	(1.6 x 17mm heat shrink
	0032310

(roll)

	
	
	
	1
	(1.6 x 19mm heat shrink
	0032310

(roll)

	
	
	
	1
	(1.6 x 8mm heat shrink
	0032310

(roll)


ASSEMBLY OPERATIONS

1.
Pre Heat soldering iron temperature to 240oc.

2.
Collect all required parts and equipment listed above.

3.
Cut a 3.1 metre length of standard 6-core cable. Shown below.

                           Fig 1.  
[image: image13.wmf]1.2 Cable Cross Section

1.0 Mechanical Specifications

1.1 Cable Schematic

Black

White

Blue

Inner Shield

Orange

Yellow

Red

Outer Shield

Outer Jacket

Outer Shield

Red,Orange, Yellow, Blue

Inner Shield

Black, White Twisted Pair

Twisted Pair

Paper/ nylon packing to make round


D-Type female 9-pin side:
Fig 2.1 
[image: image14.wmf]Pin Insert Rear View

1

Connector Plug

(9) Outer 

shield (os)

(9) Inner 

Shield (is) 

(1) Red

(6) Yellow

(8) Blue

(2) Orange

Pins

(4) Black

(5) White

Cable Grip

Outer Casing

6

7

8

9

4

5

2

3


1.
Feed outer casing, cable grip and (6 x 10mm heat shrink (black) over end of cable.

2.
Strip 20mm off outer jacket of cable to reveal coloured wires, outer shield, and nylon/paper wire packing.

3.
Cut all packing to the base.

4.
Strip 20mm off inner jacket to reveal black and white wires and the inner shield.

5.
Twist outer and inner shields together.

6.
Trim ends of wires and shields to the same length.

7.
Strip jacket of every wire 2mm to reveal copper core.

8.
Heat (1.6 x 17mm heat shrink over twisted inner and outer shields to cover naked wire, and solder end to the rear of one pin.

9.
Apply small amount of solder to ends of each wire and solder each of the 5 wires to the rear separate pins.

10.
Clamp cable grip approximately 2mm from end of outer jacket.

11.
Heat (6 x 10mm heat shrink firmly over cable grip and beginning of wires.

12.
Insert pins into correct locations (as shown in fig 2.1) and push firmly into place.

13.
Push outer casing over cable grip and wires to fit around the pin housing.

DIN male 8-pin side:

                  Fig 2.2  
[image: image15.wmf](4) Orange

Barrell

Cable Grip

(cable grip) Blue

(R1) Yellow

Connector

(5) 

Combined 

outer and 

inner 

shields 

(is+ os) 

(2)    R1 = 22.1

(6, 8) R2 = 47.5

W

W

K

(7) White

(1) Red

(3) Black

2

R1

8

3

5

7

4

1

6

R2

Connector Rear View


1.
Trim legs of (R2) 47.5 k( resistor to 4mm and 15mm.

2.
Cover resistor (R2) with (1.6 x 19mm heat shrink leaving 2mm of wire showing either end, and heat to cover extra naked wire.

3.
Form 15m side around so both resistor legs will fit into correct locations in the connector then solder into place.

4.
Trim both legs of (R1) 22.1 k( resistor to 4mm, and solder one end to pin no.2 in the rear of the connector.

5.
Feed (6 x 43mm (clear) heat shrink, barrel and (6 x 40mm heat shrink over the end of the cable.

6.
Strip 20mm off outer jacket of wire to reveal coloured wires, outer shield, and nylon/paper wire packing.

7.
Cut all packing to the base.

8.
Strip 20mm off inner jacket to reveal black and white wires and the inner shield.

9.
Twist outer and inner shields together.

10.
Heat (1.6 x 17mm heat shrink over shields to cover excess naked wire.

11.
Trim ends of wires and shields to the same length.

12.
Strip jacket of every wire 2mm to reveal copper core.

13.
Solder all wires except yellow and blue, to the correct locations on the rear of the connector.

14.
Insert (1.6 x 8 mm heat shrink over the yellow wire, solder wire to remaining leg of (R1) and then heat the heat shrink over the resistor and wire connection.

15.
Heat (6 x 40mm heat shrink over the end of the cable outer jacket to hold the beginning of the wires firmly.

16.
Attach the cable clip to the connector, and clamp the other end firmly to the cable.

17.
Solder remaining blue wire to the cable grip.

18.
Push barrel over the connector and screw the barrel firmly to the connector.

TESTING 

1.
Attach DIN male 8-pin side to the test box connector marked ‘F’.

2.
Attach D-type male 9-pin side to a Datascope finger probe.

3.
Check display is showing correct characteristics as shown below. (At correct switch positions)

       [image: image16.jpg]



               [image: image17.jpg]


                 [image: image18.jpg]R -

—— TN m—— T~

e

A AR e




      Position 2.



  Position 3.


            Position 4. 
4.
‘Play’ with wire at connections to see if any change in the display (i.e. flickering etc).

5.
If there is any movement of signal, the extension wire must be taken apart and all connections checked and re-soldered. Then tested again until results are satisfactory.

6.
Check the cable is of correct quality standard. (See VM/COP/30.11 for details).

7.
Connect DIN male 8-pin side to a Datascope monitor and attach probe on finger to check SpO2 level. (Ideal reading 95-100.)

Labelling

1.
Labels: to be attached facing upwards as looking at the top of the probe.


· 1 x CE Label

· 1 x serial no. Label

· 1x White Datascope compatible dumbbell Label.


[image: image19.wmf]White dumbell 

labell

For use on

Datascope

Compatible

Pulse Oximeters

0086

CE

Ø6 x 43 clear heat shrink

Serial No.


Quality Assurance (QA)

1.
Attach DIN male 8-pin side to the test box connector marked ‘F’.

2.
Attach D-type male 9-pin side to a Datascope finger probe.

3.
Check display is showing correct characteristics as shown below. (At correct switch positions)

       [image: image20.jpg]



               [image: image21.jpg]


                 [image: image22.jpg]R -

—— TN m—— T~

e

A AR e




      Position 2.



  Position 3.


            Position 4. 
4.
‘Play’ with wire at connections to see if any change in the display (i.e. flickering etc).

5.
If there is any movement of signal, the extension wire must be taken apart and all connections checked and re-soldered. Then tested again until results are satisfactory.

6.
Check the cable is of correct quality standard. (See VM/COP/30.11 for details).

7.
Connect DIN male 8-pin side to a Datascope monitor and attach probe on finger to check SpO2 level. (Ideal reading 95-100.)

8.
Fill and sign attached paperwork.

9.
Test 10 % of batch on DL3000 simulator.

10.
Log all results on compatibility sheet.

Packaging 

1.
Visually check all labels are attached properly

2.
Using a twist tie (bunny clip) wrap the cable and place in a small blue Viamed plastic box, ensuring the cable is inserted in a neat and tidy presentable manor.

3.
Place a serial number sticker (supplied with the batch) on the front face of the box.

4.
Place a packed and tested sticker (also containing initials of the individual who is packing) on the right hand side top left corner of the box. Do not close box.

Final QA

1.
Final inspection. Visually ensure cable sit neatly within the box and is in a presentable state.

2.
Boxes are ready to stock in stores.
















Page 5 of 5

[image: image23.png]_1070093120.dwg

_1070186341.dwg

_1071053604.dwg

_1097914526.dwg

_1070192487.dwg

_1070454884.dwg

_1070186040.dwg

_1070186221.dwg

_1070093428.dwg

_1069494359.dwg

_1069744209.dwg

_1069766336.dwg

_1069504476.dwg

_1069494204.dwg

