[image: image24.png]


 

                      SpO2 Extension Cable Assembly

	COMPANY OPERATING PROCEDURES


0019650   
     CRITIKON & SENSORMEDICS   
P965E10

VM3/COP/33.11
	Date:
  18-Dec-01
	Revision date: 6-Nov-03
	Issue:   3




[image: image1.wmf]3m

Ø

6 x 43 Clear 

heatshrink. For serial 

no.

Enlarged front view

Compushield Female 6 Pin

Thorkom Female 12 Pin











[image: image2.jpg]


                                                                           [image: image3.jpg]



Equipment required:
Soldering iron (0060120), solder (0050012), Wire stripper (0060030), Flush Cutter (0060010), Snipe nose pliers (0060021), ‘helping hand’ (0060145), Heat gun (0060100).

Parts list:
Kit and parts required. 

	Compushield female

6-pin side
	Thorkom female

12-pin side

	Qty
	Description
	Part No.
	Qty
	Description
	Part No.

	
	
	
	
	
	

	1
	Compushield female

6-pin kit
	0010712
	1
	Thorkon female 12-pin kit
	0010781

	(1)
	
[image: image4.wmf]      Housing
	kit
	
	
[image: image5.wmf]           Rear nut
	Kit

	(1)
	
[image: image6.wmf]          Connector
	kit
	
	
[image: image7.wmf]      Strain relief
	Kit

	(1)
	
[image: image8.wmf]                 Cable grip
	kit
	
	
[image: image9.wmf]               Collett 
	Kit

	(1)
	
[image: image10.wmf]        Rear insert
	kit
	
	
[image: image11.wmf]                  Pins
	Kit

	1
	
[image: image12.wmf]      Strain relief
	0010618

(50 off)
	
	
[image: image13.wmf]               Pin housing
	kit

	
	
	
	1
	(6 x 43mm Clear heat shrink
	0032331

(roll)

	
	
	
	1
	(6 x 10mm heat shrink
	0032321

(roll)

	
	
	
	2
	(1.6 x 17mm heat shrink
	0032310

(roll)


ASSEMBLY OPERATIONS

1.
Pre Heat soldering iron temperature to 240oc.

2.
Collect all required parts and equipment listed above.

3.
Cut a 3.1 metre length of special 6-core cable. Shown below.

                           Fig 1.  
[image: image14.wmf]1.2 Cable Cross Section

1.0 Mechanical Specifications

1.1 Cable Schematic

Black

White

Blue

Inner Shield

Orange

Yellow

Red

Outer Shield

Outer Jacket

Outer Shield

Red,Orange, Yellow, Blue

Inner Shield

Black, White Twisted Pair

Twisted Pair

Paper/ nylon packing to make round


Compushield female 6-pin side:

Fig 2.1 
[image: image15.wmf]Cable Grip

7

(7) 

Combined 

outer and 

inner 

shields 

(is+ os) 

(6) Red

(5) Orange

(3) Blue

(4) Yellow

(1) White

(2) Black

Rear Insert

Modify Strain Relief

Strain Relief

Connector

Housing

Front view of rear insert

1

2

3

4

5

6


1.
Cut a flat side on the strain relief as shown in fig 2.1.

2.
Feed rear insert, strain relief and (6 x 10mm heat shrink (black) over end of cable.

3.
Strip 20mm off outer jacket of cable to reveal coloured wires, outer shield, and nylon/paper wire packing.

4.
Cut all packing to the base.

5.
Strip 20mm off inner jacket to reveal black and white wires and the inner shield.


6.
Twist inner and outer shields together, and form wire backwards along cable outer jacket surface.

7.
Clamp cable grip over shields and 2mm from the end of the cable outer jacket (see fig 2.1).

8.
Push strain relief up to cable grip, and push rear insert around strain relief as far as possible.

9.
Push wires firmly into correct locations. (Apply small amount of glue if necessary). 

10.
Trim any excess wire ‘overhanging’ from the front end of the rear insert.

11.
Solder shields to the metal plate in the rear insert and trim remaining wire.

12.
Push connector onto rear insert, and use pliers to push pins down through wires. 

13.
Push housing over connector and rear insert to cover the unit.

Thorkom female 12-pin side:

Fig 2.2  
[image: image16.wmf]Cable Grip

Pin housing

(12 +  7) 

Combined 

outer and 

inner 

shields 

(is+ os) 

(1) Black

(2) White

(6) Blue

(11) Yellow

(4) Orange

(3) Red

Front view

Rear view

Pin Housing Views

1

2

3

4

5

11

12

10

9

7

8

6

Top

Top

Rear Nut

Strain Relief

Collett

Pins


1.
Feed (6 x 43mm Clear heat shrink, rear nut, collett, strain relief, cable grip, and (6 x 10mm heat shrink over the end of the cable.

2.
Strip 20mm off outer jacket of cable to reveal coloured wires, outer shield, and nylon/paper wire packing.

3.
Cut all packing to the base and strip 20mm off inner jacket of wire to reveal black and white wires and the inner shield.

4.
Combine outer and inner shields together then split to form 2 wires, each containing inner and outer shield wire.

6.
Trim ends of wires and shields to the same length.

7.
Strip jacket of every wire 2mm to reveal copper core.

8.
Place (1.6 x 17mm heat shrink over each twisted inner and outer shields to cover naked wire, and solder the ends to the rear of 2 separate pins.

9.
Solder all remaining wires to the rear of separate pins and insert firmly into correct locations shown in fig 2.2.

10.
Clamp cable grip approximately 2mm from the end of the outer jacket.

11.
Place (6 x 10mm heat shrink over cable grip and beginning of wires and heat to shrink firmly over.

12.
Push Strain relief up to the cable grip, collett over the strain relief, and screw the rear nut to the pin housing.

TESTING 

1.
Attach Thorkom female 9-pin side to a Thorkom adapter cable and then to the test box connector marked ‘J’.

2.
Attach Compushield female 6-pin to a  Sensormedics finger probe.

3.
Check display is showing correct characteristics as shown below. (At correct switch positions)

[image: image17.jpg]



[image: image18.jpg]



[image: image19.jpg]




        Pos 2. LED


   Pos 3. IR

              Pos 4. Detector

5.
‘Play’ with wire at connections to see if any change in the display (i.e. flickering etc).

6.
If there is any movement of signal, the extension wire must be taken apart and all connections checked and re-soldered. Then tested again until results are satisfactory.

7.
Check the cable is of correct quality standard. (See VM/COP/30.11 for details).

8.
Connect Thorkom female 9-pin side to a Sensormedics oxyshuttle monitor and attach probe on finger to check SpO2 level. (Ideal reading 95-100.)

Labelling

1.
Labels: to be attached facing upwards as looking at the top of the probe.


· 1 x CE Label

· 1 x serial no. Label

· 1x Orange ‘Do Not Throw Away’ Label (correct one of two is dependant of country unit is being sold to).


[image: image20.wmf]Orange dumbell 

labell

0086

CE

Ø6 x 43 clear heat shrink

Serial No.

Do Not

Throw

Away


Quality Assurance (QA)

1.
Attach Thorkom female 9-pin side to a Thorkom adapter cable and then to the test box connector marked ‘J’.

2.
Attach Compushield female 6-pin to a  Sensormedics finger probe.

3.
Check display is showing correct characteristics as shown below. (At correct switch positions)

[image: image21.jpg]



[image: image22.jpg]



[image: image23.jpg]




        Pos 2. LED


   Pos 3. IR

              Pos 4. Detector

5.
‘Play’ with wire at connections to see if any change in the display (i.e. flickering etc).

6.
If there is any movement of signal, the extension wire must be taken apart and all connections checked and re-soldered. Then tested again until results are satisfactory.

7.
Check the cable is of correct quality standard. (See VM/COP/30.11 for details).

8.
Connect Thorkom female 9-pin side to a Sensormedics oxyshuttle monitor and attach probe on finger to check SpO2 level. (Ideal reading 95-100.)

9.
Fill and sign attached paperwork.

10.
Test 10 % of batch on DL3000 simulator.

11.
Log all results on compatibility sheet.

Packaging 

1.
Visually check all labels are attached properly

2.
Using a twist tie (bunny clip) wrap the cable and place in a small blue Viamed plastic box, ensuring the cable is inserted in a neat and tidy presentable manor.

3.
Place a serial number sticker (supplied with the batch) on the front face of the box.

4.
Place a packed and tested sticker (also containing initials of the individual who is packing) on the right hand side top left corner of the box. Do not close box.

Final QA

1.
Final inspection. Visually ensure cable sit neatly within the box and is in a presentable state.

2.
Boxes are ready to stock in stores.
















Page 1 of 6

[image: image24.png]_1070190772.dwg

_1070280749.dwg

_1071055163.dwg

_1075717791.dwg

_1070280798.dwg

_1070284349.dwg

_1070280624.dwg

_1070280684.dwg

_1070280491.dwg

_1070189473.dwg

_1070189668.dwg

_1070189779.dwg

_1070189614.dwg

_1070189423.dwg

_1069744209.dwg

