[image: image1.png]


[image: image2.jpg]ARAB HEALTH
2002

Summary of products

Pulse Oximetry Sensors Viamed  Dolphin  Aristo
Finger sensors v v v
Multi-site 'Y’ sensors v v v
Ear sensors v
Extension cables v v v
Adapter cables v v v
D.O.T. Sensors v
Wrap sensors v
Disposables v v ~

Pulse oximetry sensor repair services
Pulse oximeters - Stand alone, hand held, compact PC
Pulse oximetry simulators and testers

Oxygen sensors
Oxygen monitors and analysers

Oxygen hoods
Patient monitors
Infant resuscitation cabinets

Infant resuscitators
Radiant warmers

Phototherapy masks
Phototherapy light shields

ECG cables

Temperature probes

Supra-maximal nerve stimulator

Chart & recorder paper [
VIARMED

Imail infeid





                       

MICROSTIM DBS_Mk3 – CE FILE 

Risk Assessment iaw EN ISO 14971:2000 Annex D: Possible hazards with medical devices.

	Ref.
	Hazard.

	Related part / Component posing risk.
	Sev

of

Haz.
	Like

of

Haz.
	Risk.
	Solution.
	Document referenced.
	Sev

of

Haz.
	Like

0f

Haz.
	Risk.

	D.2
	Energy hazards and contributory factors

	D.2.1
	Electricity
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.2
	Heat
	If battery polarity reversed
	3
	1
	3
	Protection diode added in circuit
	(YZ) Design File
	1
	1
	1

	D.2.3
	Mechanical force
	N/A
	2
	1
	2
	
	
	2
	1
	2

	D.2.4
	Ionising radiation
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.5
	Non ionising radiation
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.6
	Moving parts
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.7
	Unintended motion
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.8
	Suspended masses
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.9
	Patient support failure
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.10
	Pressure (vessel rupture)
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.11
	Acoustic pressure
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.12
	Vibration
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.13
	Magnetic fields 

(eg. MRI)
	N/A
	2
	2
	4
	
	F. User Instructions
	2
	1
	2

	D.3
	Biological hazards and contributory factors

	D.3.1
	Bio-contamination
	Instrument & Leads
	2
	2
	4
	Easy to clean – cleaning instructions in user instructions
	F. User Instructions
	2
	1
	2

	D.3.2
	Bio-incompatibility
	Electrodes
	1
	1
	1
	
	Manufacturers Data
	1
	1
	1

	D.3.3
	Incorrect formulation (chemical composition)
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.3.4
	Toxicity
	N/A
	1
	1
	1
	
	Manufacturers Data
	1
	1
	1

	D.3.5
	Allergenicity
	N/A
	1
	1
	1
	
	Manufacturers Data
	1
	1
	1

	D.3.6
	Mutagenicity
	N/A
	1
	1
	1
	
	Manufacturers Data
	1
	1
	1

	D.3.7
	Oncogenicity
	N/A
	1
	1
	1
	
	Manufacturers Data
	1
	1
	1

	D.3.8
	Carcinogenicity
	N/A
	1
	1
	1
	
	Manufacturers Data
	1
	1
	1

	D.3.9
	Re and/or cross infection
	Instrument & Leads
	1
	1
	1
	
	Manufacturers Data
	1
	1
	1

	D.3.10
	Pyrogenicity
	N/A
	1
	1
	1
	
	Manufacturers Data
	1
	1
	1

	D.3.11
	Inability to maintain hygienic standards
	Instrument & Leads
	2
	2
	4
	Easy to clean – cleaning instructions in user manual
	F. User Instructions
	2
	1
	2

	D.3.12
	Degradation
	Instrument & Leads
	2
	1
	2
	Care Instructions given in user manual
	F. User Instructions
	1
	1
	1

	D.4
	Environmental hazards and contributory factors

	D.4.1
	Electromagnetic fields
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.4.2
	Susceptibility to electromagnetic interference
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.4.3
	Emissions of electromagnetic interference 
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.4.4
	Inadequate supply of power 
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.4.5
	Inadequate supply of coolant
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.4.6
	Storage / operation outside prescribed environmental conditions
	Instrument
	2
	1
	2
	No specific environmental storage / operating conditions imposed
	
	2
	1
	2

	D.4.7
	Incompatibility with other devices with which the product is intended to be used
	ECG Pads
	1
	1
	1
	Standard sized ECG press studs on leads
	
	1
	1
	1

	D.4.8
	Accidental mechanical damage
	Instrument & Leads
	2
	1
	2
	Relatively robust material used. If damaged, user to assess level of damage – return for repair or spares parts available
	F. User Instructions
	1
	1
	1

	D.4.9
	Contamination due to waste products and/or device disposal
	Instrument & Leads
	1
	1
	1
	No special disposal required. Alkaline batteries used
	
	1
	1
	1

	D.5
	Hazards resulting from incorrect output of energy and substances

	D.5.1
	Electricity
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.5.2
	Radiation
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.5.3
	Volume
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.5.4
	Pressure
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.5.5
	Supply of medical gases
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.5.6
	Supply of anaesthetic agents
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.6
	Hazards related to the use of the medical device and contributory factors

	D.6.1
	Inadequate labelling
	Label
	2
	1
	2
	Product easy to use.

Label
	F. Labels
	1
	1
	1

	D.6.2
	Inadequate operating instructions
	User Instructions
	2
	1
	2
	Product easy to use.

User manual
	F. User Instructions
	1
	1
	1

	D.6.3
	Inadequate specification of accessories
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.6.4
	Inadequate specification of pre-use checks
	User Instructions / Labels
	2
	2
	4
	Product easy to use.

User manual / labels
	F. User Instructions

F. Labels
	2
	1
	2

	D.6.5
	Over-complicated operating instructions
	User Instructions
	2
	1
	2
	Product easy to use.

User manual
	F. User Instructions
	1
	1
	1

	D.6.6
	Inadequate specification of service and maintenance
	N/A
	1
	1
	1
	Instrument set up at manufacturer. Does not require regular service. Maintenance only required if performance degrades
	
	1
	1
	1

	D.6.7
	Use by unskilled / untrained personnel
	Instrument & Leads
	3
	2
	6
	Product easy to use. Warning in user manual – not for use by untrained personnel
	F. User Instructions
	2
	1
	2

	D.6.8
	Reasonable foreseeable misuse
	Instrument & Leads
	3
	2
	6
	Product easy to use. Warning in user manual – not for use by untrained personnel
	F. User Instructions
	2
	1
	2

	D.6.9
	Insufficient warning of side effects
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.6.10
	Inadequate warnings of hazards likely with re-use of single use devices
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.6.11
	Incorrect measurement and other metrological aspects
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.6.12
	Misrepresentation of results
	Instrument & Leads
	1
	1
	1
	Only for use by qualified & trained personnel – warning & in-depth explanation in user manual
	F. User Instructions
	1
	1
	1

	D.6.13
	Incompatibility with consumables / accessories / other devices
	Battery & Leads
	1
	1
	1
	Standard battery used. Leads colour coded – spares available
	F. User Instructions
	1
	1
	1

	D.6.14
	Sharp edges or points
	Instrument & Leads
	2
	2
	4
	Only if damaged, user to assess level of damage – return for repair or spares parts available
	F. User Instructions
	1
	1
	1

	D.7
	Inappropriate, inadequate or overcomplicated user interface (man/machine communication)

	D.7.1
	Mistakes & judgement errors
	N/A
	1
	1
	1
	No complex user interface
	
	1
	1
	1

	D.7.2
	Lapses and cognitive recall errors
	N/A
	1
	1
	1
	No complex user interface
	
	1
	1
	1

	D.7.3
	Slips & blunders (mental or physical)
	N/A
	1
	1
	1
	No complex user interface
	
	1
	1
	1

	D.7.4
	Violation or abbreviation of instructions,  procedures etc
	N/A
	1
	1
	1
	No complex user interface
	
	1
	1
	1

	D.7.5
	Complex or confusing control system
	N/A
	1
	1
	1
	No complex user interface
	
	1
	1
	1

	D.7.6
	Ambiguous or unclear device state
	N/A
	1
	1
	1
	No complex user interface
	
	1
	1
	1

	D.7.7
	Ambiguous or unclear presentation of settings, measurement, or other information
	N/A
	1
	1
	1
	No complex user interface
	
	1
	1
	1

	D.7.8
	Misrepresentation of results
	N/A
	1
	1
	1
	No results presented
	
	1
	1
	1

	D.7.9
	Insufficient visibility, audibility or tactility
	N/A
	1
	1
	1
	No complex user interface
	
	1
	1
	1

	D.7.10
	Poor mapping of controls to action or of displayed information to actual state
	N/A
	1
	1
	1
	No complex user interface
	
	1
	1
	1

	D.7.11
	Controversial modes or mappings as compared to existing equipment
	N/A
	1
	1
	1
	No complex user interface
	
	1
	1
	1

	D.8
	Hazards arising from function failure, maintenance and ageing and contributory factors

	D.8.1
	Erroneous data transfer
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.8.2
	Lack of, or inadequate specification for maintenance including post maintenance functional tests
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.8.4
	Inadequate maintenance
	N/A
	1
	1
	1
	Instrument set up at manufacturer. Does not require regular service. Maintenance only required if performance degrades
	
	1
	1
	1

	D.8.5
	Lack of adequate determination of end of device life
	Instrument
	1
	1
	1
	User decision based on damage or cost effectiveness of repair
	
	1
	1
	1

	D.8.6
	Loss of electrical integrity
	Instrument
	2
	2
	4
	Failure of any electrical component will result in loss of audio or visual output indication
	
	1
	1
	1

	D.8.7
	Loss of mechanical integrity
	Instrument
	1
	1
	1
	User decision based on damage or cost effectiveness of repair
	
	1
	1
	1

	D.8.8
	Inadequate packaging (contamination and / or deterioration of the device)
	Instrument & Leads
	2
	2
	4
	Boxed in protective plastic case. Bubble wrap / cardboard box through mail / courier
	M. Packaging
	1
	1
	1

	D.8.9
	Re-use and/or improper re-use
	Instrument & Leads
	1
	1
	1
	User decision on suitability for next use based on pre-use check & tolerable damage to casing
	
	1
	1
	1

	D.8.10
	Deterioration in function (gradual occlusion of fluid / gas path or change in resistance to flow, electrical conductivity) as a result of repeated use
	Instrument & Leads
	1
	1
	1
	Degradation of electrical components will eventually result in loss of audio or visual output indication
	
	1
	1
	1



T:\Kevin Rush\CE Files\CE files current\Microstim_DBS_Mk3\CE File_Microstim_DBS_Mark3\E. Risk Analysis\EN 14971 Annexe D\E. Risk Management BS14971 Annex D_Microstim.doc
    Page 1







02/06/2004
Severity of hazard: 1 = Insignificant, 2 =  Tolerable, 3  = Critical, 4  =  Intolerable.  Probability of event: 1 = Improbable, 2 = Occasional, 3 = Likely, 4 = Highly likely.  

Risk calculated as severity of hazard x probability of event, 1 – 16.  Further explanation of risk management policy – see Risk Management Policy & Definitions.


