[image: image28.png]


 

                      SpO2 Extension Cable Assembly

	COMPANY OPERATING PROCEDURES


0019733
DATEX
P973E33

VM3/COP/33.18
	Date:
  21-Dec-01
	Revision date: 19-Sep-03
	Issue:   3




[image: image1.wmf]Key-way Allignment

Enlarged front view of connector

DPK1 Female 10 Pin

10m

D-Type Female 9 Pin

Ø

6 x 43 Clear 

heatshrink. For serial 

no.


[image: image2.jpg]


                                                         [image: image3.jpg]



Equipment required:
Soldering iron (0060120), solder (0050012), Wire stripper (0060030), Flush Cutter (0060010), Snipe nose pliers (0060021), ‘helping hand’ (0060145), Heat gun (0060100).

Parts list:
Kit and parts required. (Continued over page)

	D-Type female 

9-pin Side
	DPK-1 female 

10-pin Side

	Qty
	Description
	Part No.
	Qty
	Description
	Part No.

	
	
	
	
	
	

	1
	D-type female 9-pin kit
	0010760
	1
	DPK-1 female 10-pin kit
	0010706

	(1)
	
[image: image4.wmf]  Outer Casing
	kit
	(1)
	
[image: image5.wmf]       Rubber housing
	kit

	(1)
	       
[image: image6.wmf]              Cable grip
	kit
	(1)
	       
[image: image7.wmf]              Cable grip
	kit

	(1)
	
[image: image8.wmf]           Pin Housing
	kit
	(1)
	      
[image: image9.wmf]                   Collar
	kit

	(9)
	
[image: image10.wmf]           Pins
	kit
	(10)
	
[image: image11.wmf]            Pins
	kit

	1
	1.6m 6-core cable
	0030513

(roll)
	(1)
	 
[image: image12.wmf]        Upper casing
	kit

	1
	(1.6 x 17mm heat shrink
	0032310

(roll)
	(1)
	  
[image: image13.wmf]           Cable clamp
	kit

	1
	(6 x 10mm heat shrink
	0032321

(roll)
	(1)
	    
[image: image14.wmf]      (grey) Pin housing
	kit

	
	
	
	(2)
	      
[image: image15.wmf]                  Screws
	kit

	
	
	
	(1)
	
[image: image16.wmf]            Lower casing
	kit

	
	
	
	(1)
	      
[image: image17.wmf]                 Grey Coller
	kit

	
	
	
	1
	
[image: image18.wmf]R1

               20 k( Resistor
	0032090

	
	
	
	1
	
[image: image19.wmf]             150 (F Capacitor
	0032250

	
	
	
	1
	(6 x 43mm Clear heat shrink
	0032331

(roll)

	
	
	
	1
	(1.6 x 20mm heat shrink
	0032310

(roll)

	
	
	
	1
	(6 x 25mm heat shrink
	0032321

(roll)

	
	
	
	
	
	


ASSEMBLY OPERATIONS

1.
Pre Heat soldering iron temperature to 240oc.

2.
Collect all required parts and equipment listed above.

3.
Cut a 1.6 metre length of standard 6-core cable. Shown below.

                           Fig 1.  
[image: image20.wmf]1.2 Cable Cross Section

1.0 Mechanical Specifications

1.1 Cable Schematic

Black

White

Blue

Inner Shield

Orange

Yellow

Red

Outer Shield

Outer Jacket

Outer Shield

Red,Orange, Yellow, Blue

Inner Shield

Black, White Twisted Pair

Twisted Pair

Paper/ nylon packing to make round


D-Type female 9-pin side:

Fig 2.1 
[image: image21.wmf]Outer Casing

(3) Yellow

(6) Blue

Cable Grip

9

7

6

8

Pins

Connector Plug

(9) Black

(7) Red

(1) Orange

(5) White

Pin Insert Rear View

1

2

3

4

5

(8) 

Combined 

outer and 

inner 

shields 

(is+ os) 


1.
Feed outer casing, cable grip and (6 x 10mm heat shrink (black) over end of wire.

2.
Strip 20mm off outer jacket of wire to reveal coloured wires, outer shield, and nylon/paper wire packing.

3.
Cut all packing to the base.

4.
Strip 20mm off inner jacket to reveal black and white wires and the inner shield.

5.
Twist outer and inner shields together.

6.
Trim ends of wires and shields to the same length.

7.
Strip jacket of every wire 2mm to reveal copper core.

8.
Heat (1.6 x 17mm heat shrink over twisted inner and outer shields to cover naked wire, and solder end to the rear of one pin.

9.
Clamp cable grip approximately 2mm from end of outer jacket.

10.
Place (6 x 10mm heat shrink over cable grip and beginning of wires and heat to shrink firmly over.

11.
Insert pins into correct locations (as shown in fig 2.1) and push firmly into place.

12.
Push outer casing over cable grip and wires to fit around the pin housing.

DPK-1 female 10-pin side:

                     Fig 2.2  
[image: image22.wmf]Rear View of Pin Housing

Bottom Casing

p

F

(1,8) C1 = 150

Cable clamp

Grey Collar

Pin Housing

Top Casing

(4) 

Combined 

outer and 

inner 

shields 

(is+ os) 

Rubber Housing

(5) Orange

(8) White

(6) Red

(9) Black

(7) Yellow

(4) Blue

Cable Grip

Pins

Collar

5

8

C1

R1

2

1

4

7

0

3

6

9

(1,8) R1 = ----

K

W


1.
Feed (6 x 43mm (clear) heat shrink, rubber housing, collar, cable grip and (6 x 25mm (black) heat shrink over the end of the cable.

2.
Strip 25mm off outer jacket of wire to reveal coloured wires, outer shield, and nylon/paper wire packing.

3.
Cut all packing to the base.

4.
Strip 25mm off inner jacket to reveal black and white wires and the inner shield.

5.
Strip jacket of the blue wire 5 mm and twist together with both inner and outer shields (apply small amount of solder to hold together). Heat (1.6 x 20mm heat shrink over combined shields and blue wire.

6.
Trim ends of wires to the same length.

7.
Strip jacket of every wire 2mm to reveal copper core, and solder all except white wire to the rear of individual pins.

8.
Cut each leg of the resistor to 15mm each.

9.
Cut the legs of the capacitor so the overall length is the same as the resistor.

10.
Bend legs of both resistor and capacitor to fit into the rear of two separate pins (one leg of each in one pin) then solder to pins.

11.
Insert all pins into correct locations except in pin 8 to which the white wire needs to be soldered to the same pin as that of the capacitor and resistor before inserting.

12.
Insert pin housing into the bottom casing, and screw cable clamp over cable.

13.
Push top casing onto bottom casing, and push the collar over mating cases.

14.
Push rubber housing over cases, and then the grey collar over the casing.

TESTING 

1.
Attach DPK-1 female 10-pin side to the test box connector marked ‘M’.

2.
Attach D-type female 9-pin side to a Datex finger probe.

3.
Check display is showing correct characteristics as shown below. (At correct switch positions)

[image: image23.jpg]



[image: image24.jpg]R -

—— TN m—— T~

e

A AR e




                              NB: Gap to be at bottom of display


            
            Position 2. IR, LED.
       
      Position 4. Detector

4.
If the gap is at the top of the screen then the LED is wired incorrectly.

5.
‘Play’ with wire at connections to see if any change in the display (i.e. flickering etc).

6.
If there is any movement of signal, the extension wire must be taken apart and all connections checked and re-soldered. Then tested again until results are satisfactory.


7.
Check the cable is of correct quality standard. (See VM/COP/30.11 for details).

8.
Connect female 10 pin side to the Datex monitor and attach probe on finger to check SpO2 level. (Ideal reading 95-100.)

Labelling

1.
Labels:
 to be attached facing upwards as looking at the top of the probe.


· 1 x CE Label

· 1 x Serial no. Label

· 1 x Brown Datex Dumbbell Label

[image: image25.wmf]Ø6 x 43 clear heat shrink

CE

0086

For use on

Datex

Compatible

Pulse Oximeters

Brown dumbell 

labell

Serial No.


Quality Assurance (QA)

1.
Attach DPK-1 female 10-pin side to the test box connector marked ‘M’.

2.
Attach D-type female 9-pin side to a Datex finger probe.

3.
Check display is showing correct characteristics as shown below. (At correct switch positions)

[image: image26.jpg]



[image: image27.jpg]R -

—— TN m—— T~

e

A AR e




                              NB: Gap to be at bottom of display


            
            Position 2. IR, LED.
       
      Position 4. Detector

4.
If the gap is at the top of the screen then the LED is wired incorrectly.

5.
‘Play’ with wire at connections to see if any change in the display (i.e. flickering etc).

6.
If there is any movement of signal, the extension wire must be taken apart and all connections checked and re-soldered. Then tested again until results are satisfactory.


7.
Check the cable is of correct quality standard. (See VM/COP/30.11 for details).

8.
Connect female 10 pin side to the Datex monitor and attach probe on finger to check SpO2 level. (Ideal reading 95-100.)

9.
Fill and sign attached paperwork.

10.
Test 10 % of batch on DL3000 simulator.

11.
Log all results on compatibility sheet.

Packaging 

1.
Visually check all labels are attached properly

2.
Using a twist tie (bunny clip) wrap the cable and place in a small blue Viamed plastic box, ensuring the cable is inserted in a neat and tidy presentable manor.

3.
Place a serial number sticker (supplied with the batch) on the front face of the box.

4.
Place a packed and tested sticker (also containing initials of the individual who is packing) on the right hand side top left corner of the box. Do not close box.

Final QA

1.
Final inspection. Visually ensure cable sit neatly within the box and is in a presentable state.

2.
Boxes are ready to stock in stores.
















Page 6 of 6

[image: image28.png]_1070368914.dwg

_1070369328.dwg

_1070435674.dwg

_1070456987.dwg

_1075718860.dwg

_1070435707.dwg

_1070369446.dwg

_1070369781.dwg

_1070434425.dwg

_1070369566.dwg

_1070369391.dwg

_1070369122.dwg

_1070369243.dwg

_1070368994.dwg

_1070093120.dwg

_1070093428.dwg

_1069494359.dwg

_1069744209.dwg

_1069494204.dwg

