Technical Description of the Microstim

1. Power supply

a. VR1 is a 5v voltage regulator supplied from a 9v Alkaline battery

b. During pulse delivery up to 90 volts is delivered into 1kohm impedance it is possible to pull the battery down

c. C5 :C2 & C3 are reservoir capacitors

d. The current arrangement was chosen due to the Value in mfd required and the physical size of capacitors available at the time of design

e. Values of C8 & C4 have been chosen to meet the criteria of various IC regulator manufacturers

f. C6 & C2  can be combined but must be zero leakage or they will drain the battery even when not in use. 

g. C4 is recommended to be as near to the PIC as possible

2. Function selection

a. The four switches each control one function 

b. To prevent the unit being inadvertently left switched on power is only supplied when a switch is pressed

c. R7.R8.R9.R10 are pull down resistors which enable the PIC to determine which switch has been pressed

3. Mode controller

a. Th eoscillator is a standard PIC design

b. Two small capacitors have been re-added to the crystal in line with the PIC manufacturers datasheet

4. Battery level Indicator

a. This consists of a tri state LED 

b. When the battery voltage is greater than 7.8 volts the green LED is illuminated gradually turning to Red as the voltage falls to 5 volts

5. Current Detection

a. It is necessary to ensure that good conductivity ( about 1Kohm) is experienced through the patient during activation..

b. R6 and ZD3 are used so that when sufficient current passes through ZD3 Pin 1 on the PIC goes low and generates an output on pin 8 activating and Audio indicator

c. C1  was found to be necessary for noise filtering

6. Power amplifier

a. Pulses from pin 9 on the PIC are fed via TR1 & TR2 into a step up transformer  comprising T1,T2 generating up to 90v

b. The easiest check for the microstim is to turn gain to maximum and place a mains neon across the output.

C:\Documents and Settings\manager6\My Documents\Technical Description of the Microstim.doc

