| Measurement Range: | 0 % ... 100 % oxygen (at atmospheric pressure) |
| --- | --- |
| Output in Ambient Air: | 9 to 13 mV |
| Electrical Interface: | 3.5 mm mono phone Jack |
| Repeatability: | < 1% vol. O2 @ constant temperature and pressure |
| Linearity error: | < 3% relative |
| Accuracy: | 0 % … 3.3 % oxygen: ± 0.1 % absolute 3.4 % … 100 % oxygen: ± 3 % relative |
| Response time: | < 5 s to 90 % of final value; < 40 s fall time from 20.95 % to 0.1 % oxygen (when 100 % nitrogen applied) |
| Zero Offset Voltage: | < 40 μV in 100 % nitrogen |
| Cross Interference: | < 0.1 % oxygen response to: 15 % CO2 balance N2; 10 % CO balance N2; 3000 ppm NO balance N2; 3000 ppm C3H8 balance N2; 500 ppm H2S balance N2; 500 ppm SO2 balance N2; 1000 ppm benzene balance N2 |
| Influence of Humidity: | - 0.03% rel. per % RH at 25°C |
| Pressure Range: | 0.6 bar … 2 bar (ppO2 0 … 1250 mbar O2) |
| Influence of Mechanical Shock: | < 1% relative after a fall from 1m |
| Operating Temperature: | 0°C to 50°C |
| Temperature Compensation: | Built-in NTC compensation |
| Effect of Temperature Compensation (steady state): | between +25 °C and +40 °C: 3 % relative error between 0 °C and +50 °C: 8 % relative error |
| Operating Humidity: | 0 - 99% RH non-condensing |
| Long Term Output Drift: | < 1% vol. oxygen per month |
|  | Typically < -15% relative over lifetime |
| Storage Temperature: | -20°C to +50°C |
| Recommended Storage: | +5°C to +15°C |
| Recommended Load: | ≥10 kOhms |
| Warm-Up Time: | < 30 minutes, after replacement of sensor |
| Nominal Sensor Lifetime: | ≥500,000% vol oxygen hours |
| Weight: | Approximately 28 grams |
| Warranty Period: | 15 months (including 3-month shelf life) |
| Confirmations: | PTB-A 18.10 approved |
|  | Meet BAR 97 requirements |
| Part No.: | OOA101-1V: 1002499 (REF 0110177 / R-17AVG) |
| All specifications are applicable at standard conditions: 1013 mbar, 25°C dry ambient air. |  |

Mechanical draft:

General tolerances ISO 2768-c

Color of housing parts: RAL 7040

Customized Design Specification:

Product Labeling

| OOA101-1V: 1002499 (REF 0110177 / R-17AVG) |  |
| --- | --- |
| Label on Sensor (90x15) | Label on Blister- Card (63,5x46,6) |

Customer Approval

Date:

Signature: