[image: image26.png]


 

                      SpO2 Extension Cable Assembly

	COMPANY OPERATING PROCEDURES


0019910
  CRITIKON
P991E10

VM3/COP/33.19
	Date:
  21-Dec-01
	Revision date: 06 November 2003
	Issue:   3




[image: image1.wmf]Compushield Female 6 Pin

Lemo Male 7 Pin

Enlarged front view

3m

Ø

6 x 43 Clear 

heatshrink. For serial 

no.


[image: image2.jpg]


                                                             [image: image3.jpg]



Equipment required:
Soldering iron (0060120), solder (0050012), Wire stripper (0060030), Flush Cutter (0060010), Snipe nose pliers (0060021), ‘helping hand’ (0060145), Heat gun (0060100).

Parts list:
Kit and parts required. (Continued over page)

	Compushield female

6-pin side
	Lemo male

7-pin Side

	Qty
	Description
	Part No.
	Qty
	Description
	Part No.

	
	
	
	
	
	

	1
	Compushield female

6-pin kit
	0010701
	1
	Lemo male 

7-pin kit
	0010651

	(1)
	
[image: image4.wmf]      Housing
	kit
	(1)
	
[image: image5.wmf]           Housing
	Kit

	(1)
	
[image: image6.wmf]          Connector
	kit
	(1)
	
[image: image7.wmf]           Connector
	Kit

	(1)
	
[image: image8.wmf]                 Cable grip
	kit
	(1)
	
[image: image9.wmf]      Upper sleeve
	Kit

	(1)
	
[image: image10.wmf]        Rear insert
	kit
	(1)
	
[image: image11.wmf]         Lower sleeve
	Kit

	1
	
[image: image12.wmf]      Strain relief
	0010618

(50 off)
	(1)
	
[image: image13.wmf]                 Collett
	Kit

	
	
	
	(1)
	
[image: image14.wmf]             Rear Housing
	Kit

	
	
	
	(1)
	
[image: image15.wmf]         Strain relief
	0030654

	
	
	
	1
	(6 x 43mm Clear heat shrink
	0032331

(roll)

	
	
	
	1
	(1.6 x 17mm heat shrink
	0032310

(roll)

	
	
	
	1
	(6 x 25mm heat shrink
	0032321

(roll)

	
	
	
	
	
	


ASSEMBLY OPERATIONS

1.
Pre Heat soldering iron temperature to 240oc.

2.
Collect all required parts and equipment listed above.

3.
Cut a 3.1 metre length of special 6-core cable. Shown below.

                           Fig 1.  
[image: image16.wmf]1.2 Cable Cross Section

1.0 Mechanical Specifications

1.1 Cable Schematic

Black

White

Blue

Inner Shield

Orange

Yellow

Red

Outer Shield

Outer Jacket

Outer Shield

Red,Orange, Yellow, Blue

Inner Shield

Black, White Twisted Pair

Twisted Pair

Paper/ nylon packing to make round


Compushield female 6-pin side:

Fig 2.1 
[image: image17.wmf](2) Black

(1) White

(5) Orange

(6) Red

Rear Insert

6

5

4

3

2

1

7

Rear view of rear insert

(7) 

Combined 

outer and 

inner 

shields 

(is+ os) 

Modify Strain Relief

Housing

(4) Yellow

Connector

(3) Blue

Cable Grip

Strain Relief


1.
Cut a flat side on the strain relief as shown in fig 2.1.

2.
Feed rear insert and strain relief over end of cable.

3.
Strip 20mm off outer jacket of cable to reveal coloured wires, outer shield, and nylon/paper wire packing.

4.
Cut all packing to the base.

5.
Strip 20mm off inner jacket to reveal black and white wires and the inner shield.


6.
Twist inner and outer shields together, and form wire backwards along cable outer jacket surface.

7.
Clamp cable grip over shields and 2mm from the end of the cable outer jacket (see fig 2.1).

8.
Push strain relief up to cable grip, and push rear insert around strain relief as far as possible.

9.
Push wires firmly into correct locations. (Apply small amount of glue if necessary). 

10.
Trim any excess wire ‘overhanging’ from the front end of the rear insert.

11.
Solder shields to the metal plate in the rear insert and trim remaining wire.

12.
Push connector onto rear insert, and use pliers to push pins down through wires. 

13.
Push housing over connector and rear insert to cover the unit.

Lemo male 7-pin side:

                     Fig 2.2  
[image: image18.wmf](2) 

Combined 

outer and 

inner 

shields 

(is+ os) 

All Keyways facing top

Rear Nut

Strain Relief

(3) Orange

(5) Yellow

Collett

(1) Blue

Metal Sleeving

Connector Pin

(6) Black

(7) White

(4) Red

Connector Pin Rear View

1

Housing

2

3

4

5

6

7


1.
Feed (6 x 43mm (clear) heat shrink, strain relief, rear nut, collett and (6 x 25mm (black) heat shrink over end of cable.

2.
Strip 20mm off outer jacket of cable to reveal coloured wires, outer shield, and nylon/paper wire packing.

3.
Cut all packing and blue wire to the base.

4.
Strip 20mm off inner jacket to reveal black and white wires and the inner shield.

5.
Twist outer and inner shields together.

6.
Trim ends of wires to 13mm long.

7.
Strip jacket of every wire 2mm to reveal copper core.

8
Heat (6 x 25mm (black) heat shrink up to the end of the cable.

9.
Fold shields back along the surface and glue down over the heatshrink.

10.
Solder ends of every wire and shields to the correct positions on the rear of the connector.

11.
Slide collett over shields and to the end of the cable.

12.
Fit metal sleeves over the connector, and insert into the housing ensuring all keyways lie up.

13.
Holding the housing fixed with a spanner, tighten the rear nut.

14.
Push the strain relief onto the rear nut.

TESTING 

1.
Attach Lemo male 7-Pin side to a Dynamap cable then to the test box connector marked ‘B’.

2.
Attach Compushield female 6-pin side to a Sensormedics finger probe.

3.
Check display is showing correct characteristics as shown below. (At correct switch positions)

                                                                                                                                                                                        [image: image19.jpg]



[image: image20.jpg]



[image: image21.jpg]R -

—— TN m—— T~

e

A AR e




Position 2. IR
                                                Position 3. LED                                            Position 4. Detector


5.
‘Play’ with wire at connections to see if any change in the display (i.e. flickering etc).

6.
If there is any movement of signal, the extension wire must be taken apart and all connections checked and re-soldered. Then tested again until results are satisfactory.


7.
Check the cable is of correct quality standard. (See VM/COP/30.11 for details).

Labelling

1.
Labels:
 to be attached facing upwards as looking at the top of the probe.


· 1 x CE Label

· 1 x Serial no. Label

· 1x Orange ‘Do Not Throw Away’ Label (correct one of two is dependant of country unit is being sold to).

[image: image22.wmf]Orange dumbell 

labell

Do Not

Throw

Away

0086

CE

Ø6 x 43 clear heat shrink

Serial No.


Quality Assurance (QA)

1.
Attach Lemo male 7-Pin side to a Dynamap cable then to the test box connector marked ‘B’.

2.
Attach Compushield female 6-pin side to a Sensormedics finger probe.

3.
Check display is showing correct characteristics as shown below. (At correct switch positions)                                                                                                                                                                                        [image: image23.jpg]



[image: image24.jpg]



[image: image25.jpg]R -

—— TN m—— T~

e

A AR e




Position 2. IR
                                                Position 3. LED                                            Position 4. Detector


5.
‘Play’ with wire at connections to see if any change in the display (i.e. flickering etc).

6.
If there is any movement of signal, the extension wire must be taken apart and all connections checked and re-soldered. Then tested again until results are satisfactory.


7.
Check the cable is of correct quality standard. (See VM/COP/30.11 for details).

8.
Fill and sign attached paperwork.

Packaging 

1.
Visually check all labels are attached properly

2.
Using a twist tie (bunny clip) wrap the cable and place in a small blue Viamed plastic box, ensuring the cable is inserted in a neat and tidy presentable manor.

3.
Place a serial number sticker (supplied with the batch) on the front face of the box.

4.
Place a packed and tested sticker (also containing initials of the individual who is packing) on the right hand side top left corner of the box. Do not close box.

Final QA

1.
Final inspection. Visually ensure cable sit neatly within the box and is in a presentable sate.

2.
Boxes are ready to stock in stores.
















Page 1 of 5

[image: image26.png]_1070354636.dwg

_1070355037.dwg

_1070355164.dwg

_1070446489.dwg

_1071060041.dwg

_1070447212.dwg

_1070438225.dwg

_1070355092.dwg

_1070354858.dwg

_1070354925.dwg

_1070354786.dwg

_1070189473.dwg

_1070189668.dwg

_1070189779.dwg

_1070189614.dwg

_1070189423.dwg

_1069744209.dwg

