[image: image15.png]


 

                      SpO2 Extension Cable Assembly

	COMPANY OPERATING PROCEDURES


0019928    MARQUTTE
P992E8

VM3/COP/33.20
	Date:
  21-Dec-01
	Revision date: 14-Mar-05
	Issue:   3



       
[image: image1.wmf]2.4m

P956E8 Extension cable

Ø

6 x 43 Clear 

heatshrink. For serial 

no.

SpO

2

Marquette Male 11 Pin


[image: image2.jpg]


 




   [image: image3.jpg]


                                                                   

Equipment required:
Soldering iron (0060120), solder (0050012), Wire stripper (0060030), Flush Cutter (0060010), Snipe nose pliers (0060021), ‘helping hand’ (0060145), Heat gun (0060100).

Parts list:
Kit and parts required. 

	D-Type female 

9-pin Side
	Marquette male

11-pin Side

	Qty
	Description
	Part No.
	Qty
	Description
	Part No.

	
	
	
	
	
	

	1
	D-type extension cable

(Female side)
	00956E8
	1
	Marquette male 

11-pin kit
	0010770

	
	
	
	(1)
	           
[image: image4.wmf]            Cable grip
	kit

	
	
	
	(1)
	
[image: image5.wmf]        Upper casing
	kit

	
	
	
	(1)
	
[image: image6.wmf]     Lower casing
	kit

	
	
	
	(1)
	
[image: image7.wmf]              Connector
	kit

	
	
	
	1
	(6 x 43mm Clear heat shrink
	0032331

(roll)

	
	
	
	1
	(1.6 x 22 mm heat shrink
	0032310

(roll)

	
	
	
	1
	(6 x 10mm heat shrink
	0032321

(roll)


ASSEMBLY OPERATIONS

1.
Pre Heat soldering iron temperature to 240oc.

2.
Collect all required parts and equipment listed above.

D-type female 9-pin side:

  
[image: image8.wmf]2.4m

D-Type Female 9 Pin

D-Type Male 9 Pin


 Fig 2.1.  00956E8 Extension Cable.
1.
The 00956E8 cable is already assembled. The female side does not need to be modified, but the male side needs to be cut off ready for the Marquette male 11-pin connector to be fitted keeping cable as long as possible.

Marquette male 11-pin side:
Fig 2.1 
[image: image9.wmf](6) 

Combined 

outer and 

inner 

shields 

(is+ os) 


1.
Feed (6 x 43mm (clear) heat shrink, rubber housing, (6 x 10mm (black) heat shrink and cable grip over the end of the cable.

2.
Strip 25mm off outer jacket of cable to reveal coloured wires, outer shield, and nylon/paper wire packing.

3.
Cut all packing to the base.

4.
Strip 25mm off inner jacket to reveal black and white wires and the inner shield.

5.
Twist outer and inner shields together.

6.
Strip jacket of every wire 2mm to reveal copper core.

7.
Heat (1.6x 22mm heat shrink over twisted inner and outer shields to cover excess naked wire.

8.
Take a 10mm piece of wire, strip 2mm off each end and solder one end to Pin 7 and the other to Pin 8.

9.
Solder ends of every wire and shields to the correct positions on the rear of the connector top and bottom (shown in fig 2.2.)

10.
Clamp cable grip approximately 2mm from end of outer jacket.

11.
Place (6 x 10mm heat shrink over cable grip and beginning of wires and heat to shrink firmly over.

12.
Push upper and lower casing around the connector and cable grip.

13.
Push rubber housing over cable grip and wires to fit firmly around the combined upper and lower casing.

TESTING 

1.
Use the multi meter to check continuity between every wire from pin to pin.

2.
Check the cable is of correct quality standard. (See VM/COP/30.11 for details).

   Front View




   Front View

   Female 9 Pin



   Male 11 Pin

               
[image: image10.wmf]1

2

3

4

5

6

7

8

9

                               
[image: image11.wmf]
Labelling

1.
Labels:
 to be attached facing upwards as looking at the top of the probe.


· 1 x CE Label

· 1 x Serial no. Label

· 1x Orange ‘Do Not Throw Away’ Label (correct one of two is dependant of country unit is being sold to).

[image: image12.wmf]0086

CE

Serial No.

Do Not 

Throw 

Away

Orange Dumbell

Label

Dia 6mm x 43 clear heat shrink


Quality Assurance (QA)

1.
Use the multi meter to check continuity between every wire from pin to pin.

2.
Check the cable is of correct quality standard. (See VM/COP/30.11 for details).

3.
Fill and sign attached paperwork.

4.
Test 10 % of batch on DL3000 simulator.

5.
Log all results on compatibility sheet.

   Front View




   Front View

   Female 9 Pin



   Male 11 Pin

          
[image: image13.wmf]1

2

3

4

5

6

7

8

9

                                
[image: image14.wmf]
Packaging 

1.
Visually check all labels are attached properly

2.
Using a twist tie (bunny clip) wrap the cable and place in a small blue Viamed plastic box, ensuring the cable is inserted in a neat and tidy presentable manor.

3.
Place a serial number sticker (supplied with the batch) on the front face of the box.

4.
Place a packed and tested sticker (also containing initials of the individual who is packing) on the right hand side top left corner of the box. Do not close box.

Final QA

1.
Final inspection. Visually ensure cable sit neatly within the box and is in a presentable state.

2.
Boxes are ready to stock in stores.
















Page 5 of 5

[image: image15.png]_1127117933.dwg

_1127127484.dwg

_1172307360.dwg

_1172307745.dwg

_1127119016.dwg

_1127125334.dwg

_1127118387.dwg

_1127116704.dwg

_1127117882.dwg

_1070455461.dwg

