	     [image: image1.png]



	VIAMED
	
	Viamed Limited - 15 Station Road - Cross Hills

Keighley - West Yorkshire BD20 7DT - United Kingdom
Tel: +44 1535 634542                 Fax: +44 1535 635582

Email: info@viamed.co.uk  Website: www.viamed.co.uk



[image: image1.png]
Specification Datasheet – 0110177
Viamed

R-17AVG Oxygen Sensor
	    Specifications


	Part number
	0110177

	Model
	R-17AVG

	Output in ambient air
	9 mV – 13 mV 

	Output connector/connection
	3.5 mm mono phone jack

	Input connector/connection
	M16 x 1 (5/8-24 UNEF) Intermediate thread, metric/unified extra fine

	Input O ring
	Yes

	Measurement range
	0 – 100% oxygen (at atmospheric pressure)

	Repeatability
	< 1% volume oxygen at constant temperature and pressure

	Influence of pressure
	Output proportional to change in oxygen partial pressure

	Pressure range
	0.6 to 2.0 bar

	Linearity error
	< 3% relative

	Accuracy
	0 % to 3.3 % oxygen:     +/- 0.1 % absolute

3.4 % to 100 % oxygen: +/- 3.0 % relative

	Warm up time
	< 30 minutes after sensor installed 

	Zero offset voltage
	< 40µV in 100% nitrogen

	Cross interference
	< 0.1% volume oxygen response to: 

15% CO2 balance N2, 10% CO balance N2, 3000 ppm NO balance N2, 

3000 ppm C3H8 balance N2, 500 ppm H2S balance N2, 

500 ppm SO2 balance N2, 1000 ppm benzene balance N2

	Response time (t90)
	< 5 s for 90% of final value.

< 40 s fall time from 20.95% to 0.1% oxygen, when 100% nitrogen applied.

	Operating humidity
	0 – 99% R.H. (non-condensing).

	Influence of humidity
	- 0.03% relative per % R.H. at 25 ºC.

	Influence of mechanical shock
	< 1% relative after fall from 1m

	Operating temperature range
	0 to +50 ºC

	Storage temperature
	-20 to +50 ºC

	Recommended storage temperature
	+5 to +15 ºC

	Temperature compensation
	Integrated NTC thermistor

	Temperature compensation error

(steady state)
	+25 to +40 ºC: 3% relative

    0 to +50 ºC: 8% relative

	Nominal sensor life
	≥ 500,000 % oxygen hours

	Long term output drift
	< 1% volume oxygen per month, typically < -15% over sensor lifetime.

	Recommended load
	≥ 10 K(

	Weight (approximate)
	28 g

	Flow diverter
	Not included

	Packaging
	Gas barrier bag – transparent.

	Warranty period
	12 months from date of sales invoice

	All specifications are applicable at standard conditions: 1013 mbar, 25°C dry ambient air.

Specifications subject to change.



V1.2    10/03/2020
Page 1 of 1

	Drawings


[image: image2.jpg]



Measurements are in mm

Page 2 of 4

	Images


[image: image3.jpg]



[image: image4.png]VIRMED R-17AVG Oxygen Sensor
15 Station Road « Cross Hills « M SIN V123456
8020 757  Unitod Kingdom A @ CE€  Lomn  NNUMINTINN




[image: image5.png]M VIAMED

15 Station Road * Crass Hilis « Tel: +44 (0)1535 634542
Keighley West Yorkshire + Email: info@viamed.co.uk
BD20 7DT » United Kingdom Website: www.viamed.co.uk

R-17AVG Oxygen Sensor

M 2017-11 [
SIN V123456
(NI ®

Depending on the ambisnt conditions the sensor may requive up 1o 30 minutes to reach

signal stabilty afrer installation. The sensor contains caustic Jiquid. In case of leakage
avoid contact with eyes and skin! Recommended storage temperature: 5°C - 15





Page 3 of 4

	Labelling



Oxygen sensor label 90.0mm x 15.0mm

White vinyl label, text & graphics - black

[image: image6.png]4.5

34.8
41.4

5/8-24UNEF*(M16x1)

028
@29

) (+)

Phone Jack





Gas barrier bag label 63.5mm x 46.6mm

White label, text & graphics – black

[image: image7.jpg]



Page 4 of 4
[image: image8.jpg]


