[image: image21.png]


 

                      SpO2 Extension Cable Assembly

	COMPANY OPERATING PROCEDURES


0019968
  S&W
P996E8

VM3/COP/33.24
	Date:
  28-Dec-01
	Revision date: 19-Sep-03
	Issue:   3




[image: image1.wmf]Ø

6 x 43 Clear 

heatshrink. For serial 

no.

P956E8 Extension cable

Enlarged front view

SPK1-N Female 10 pin

Key-way Allignment


[image: image2.jpg]


                                                                [image: image3.jpg]



Equipment required:
Soldering iron (0060120), solder (0050012), Wire stripper (0060030), Flush Cutter (0060010), Snipe nose pliers (0060021), ‘helping hand’ (0060145), Heat gun (0060100).

Parts list:
Kit and parts required. (Continued over page)

	D-Type female 

9-pin Side
	SPK1-N female 

10-pin Side

	Qty
	Description
	Part No.
	Qty
	Description
	Part No.

	
	
	
	
	
	

	1
	D-type extension cable

(Female side)
	00956E4
	1
	SPK1-N female 10-pin kit
	0010709

	
	
	
	(1)
	
[image: image4.wmf]       Rubber housing
	kit

	
	
	
	(1)
	       
[image: image5.wmf]              Cable grip
	kit

	
	
	
	(1)
	      
[image: image6.wmf]                   Collar
	kit

	
	
	
	(10)
	
[image: image7.wmf]            Pins
	kit

	
	
	
	(1)
	 
[image: image8.wmf]        Upper casing
	kit

	
	
	
	(1)
	  
[image: image9.wmf]          Cable clamp
	kit

	
	
	
	(1)
	    
[image: image10.wmf]              Pin housing
	kit

	
	
	
	(2)
	      
[image: image11.wmf]                  Screws
	kit

	
	
	
	(1)
	
[image: image12.wmf]            Lower casing
	kit

	
	
	
	(1)
	      
[image: image13.wmf]             Pink Collar
	kit

	
	
	
	1
	(6 x 43mm Clear heat shrink

(roll)
	0032331

	
	
	
	1
	(1.6 x 20mm heat shrink

(roll)
	0032310

	
	
	
	1
	(6 x 25mm heat shrink

(roll)
	0032321

	
	
	
	
	
	


ASSEMBLY OPERATIONS

1.
Pre Heat soldering iron temperature to 240oc.

2.
Collect all required parts and equipment listed above.

D-type female 9-pin side:


[image: image14.wmf]2.4m

D-Type Female 9 Pin

Ø

6 x 43 Clear 

heatshrink. For serial 

no.

D-Type Male 9 Pin


 Fig 2.1.  00956E8 Extension Cable.
1.
The 00956E8 cable is already assembled. The female side does not need to be modified, but the male side needs to be cut off ready for the SPK1-N female 10-pin connector to be fitted.


SPK1-N female 10-pin side:

Fig 2.2  
[image: image15.wmf](7) 

Combined 

outer and 

inner 

shields 

(is+ os) 

Cable grip

Rubber Housing

Collar

Bottom Casing

Pin Housing

Top Casing

Colour Collar

Rear View of Pin Housing

9

1

2

5

4

0

8

7

3

6

(8) Blue

(9) Yellow

Pins

(2) Black

(6) Orange

(0) Red

(1) White


1.
Feed (6 x 43mm (clear) heat shrink, rubber housing, collar, cable grip and (6 x 25mm (black) heat shrink over the end of the cable.

2.
Strip 25mm off outer jacket of cable to reveal coloured wires, outer shield, and nylon/paper wire packing.

3.
Cut all packing to the base.

4.
Strip 25mm off inner jacket to reveal black and white wires and the inner shield.

5.
Twist inner and outer shields together, and heat (6 x 25mm heat shrink over excess naked wire.

6.
Trim ends of wires to the same length.

7.
Strip jacket of every wire 2mm to reveal copper core, and solder to the rear of individual pins.

8.
Insert all pins into correct locations.

9.
Insert pin housing into the bottom casing, and screw cable clamp over cable.

10.
Push top casing onto bottom casing, and push the collar over mating cases.

11.
Push rubber housing over cases, and then the pink collar over the front of the casing.

TESTING 

1.
Attach SPK1-N female 10-pin side to the S&W test box and then to the test box connector marked ‘A’.

2.
Attach D-type female 9-pin side to a Nellcor finger probe.

3.
Check display is showing correct characteristics as shown below. (At correct switch positions)

[image: image16.jpg]



[image: image17.jpg]R -

—— TN m—— T~

e

A AR e




                              NB: Gap to be at bottom of display


            
            Position 1. IR, LED.
       
      Position 4. Detector

4.
If the gap is at the top of the screen then the LED is wired incorrectly.

5.
‘Play’ with wire at connections to see if any change in the display (i.e. flickering etc).

6.
If there is any movement of signal, the extension wire must be taken apart and all connections checked and re-soldered. Then tested again until results are satisfactory.


7.
Check the cable is of correct quality standard. (See VM/COP/30.11 for details).

8.
Connect male 9 pin side to the Nellcor monitor and attach probe on finger to check SpO2 level. (Ideal reading 95-100.)

Labelling

1.
Labels:
 to be attached facing upwards as looking at the top of the probe.


· 1 x CE Label

· 1 x Serial no. Label

· 1x Orange ‘Do Not Throw Away’ Label (correct one of two is dependant of country unit is being sold to).

[image: image18.wmf]Orange dumbell 

labell

Do Not

Throw

Away

0086

CE

Ø6 x 43 clear heat shrink

Serial No.


Quality Assurance (QA)

1.
Attach SPK1-N female 10-pin side to the S&W test box and then to the test box connector marked ‘A’.

2.
Attach D-type female 9-pin side to a Nellcor finger probe.

3.
Check display is showing correct characteristics as shown below. (At correct switch positions)

[image: image19.jpg]



[image: image20.jpg]R -

—— TN m—— T~

e

A AR e




                              NB: Gap to be at bottom of display


            
            Position 1. IR, LED.
       
      Position 4. Detector

4.
If the gap is at the top of the screen then the LED is wired incorrectly.

5.
‘Play’ with wire at connections to see if any change in the display (i.e. flickering etc).

6.
If there is any movement of signal, the extension wire must be taken apart and all connections checked and re-soldered. Then tested again until results are satisfactory.


7.
Check the cable is of correct quality standard. (See VM/COP/30.11 for details).

8.
Connect male 9 pin side to the Nellcor monitor and attach probe on finger to check SpO2 level. (Ideal reading 95-100.)

9.
Fill and sign attached paperwork.

Packaging 

1.
Visually check all labels are attached properly

2.
Using a twist tie (bunny clip) wrap the cable and place in a small blue Viamed plastic box, ensuring the cable is inserted in a neat and tidy presentable manor.

3.
Place a serial number sticker (supplied with the batch) on the front face of the box.

4.
Place a packed and tested sticker (also containing initials of the individual who is packing) on the right hand side top left corner of the box. Do not close box.

Final QA

1.
Final inspection. Visually ensure cable sit neatly within the box and is in a presentable state.

2.
Boxes are ready to stock in stores.
















Page 5 of 5

[image: image21.png]_1070369122.dwg

_1070459029.dwg

_1070459281.dwg

_1071038473.dwg

_1070459081.dwg

_1070369328.dwg

_1070369391.dwg

_1070458258.dwg

_1070369243.dwg

_1070368914.dwg

_1070368994.dwg

_1069494359.dwg

_1070101508.dwg

_1069494204.dwg

