[image: image14.png]


 

                      SpO2 Extension Cable Assembly

	COMPANY OPERATING PROCEDURES


0019988
  DRAEGER
P998E8

VM3/COP/33.26
	Date:
  28-Dec-01
	Revision date: 19-Sep-03
	Issue:   3




[image: image1.wmf]Ø

6 x 43 Clear 

heatshrink. For serial 

no.

P956E8 Extension cable

Draeger Male 8 Pin

Enlarged front view


[image: image2.jpg]


                                                             [image: image3.jpg]



Equipment required:
Soldering iron (0060120), solder (0050012), Wire stripper (0060030), Flush Cutter (0060010), Snipe nose pliers (0060021), ‘helping hand’ (0060145), Heat gun (0060100).

Parts list:
Kit and parts required. (Continued over page)

	D-Type female 

9-pin Side
	Draeger male

8-pin Side

	Qty
	Description
	Part No.
	Qty
	Description
	Part No.

	
	
	
	
	
	

	1
	D-type extension cable

(Female side)
	00956E8
	1
	Draeger male 

8-pin kit
	0030657

	
	
	
	(1)
	
[image: image4.wmf]       Housing
	Kit

	
	
	
	(1)
	
[image: image5.wmf]           Connector
	Kit

	
	
	
	(1)
	
[image: image6.wmf]      Upper sleeve
	Kit

	
	
	
	(1)
	
[image: image7.wmf]         Lower sleeve
	Kit

	
	
	
	(1)
	
[image: image8.wmf]                 Collett
	Kit

	
	
	
	(1)
	
[image: image9.wmf]             Rear Housing
	Kit

	
	
	
	(1)
	
[image: image10.wmf]         Strain relief
	Kit

	
	
	
	1
	(6 x 43mm Clear heat shrink
	0032331

(roll)

	
	
	
	1
	(1.6 x 17mm heat shrink
	0032310

(roll)

	
	
	
	1
	(6 x 25mm heat shrink
	0032321

(roll)


ASSEMBLY OPERATIONS

1.
Pre Heat soldering iron temperature to 240oc.

2.
Collect all required parts and equipment listed above.

D-type female 9-pin side:


[image: image11.wmf]2.4m

D-Type Female 9 Pin

Ø

6 x 43 Clear 

heatshrink. For serial 

no.

D-Type Male 9 Pin


 Fig 2.1.  00956E8 Extension Cable.

1.
The 00956E8 cable is already assembled. The female side does not need to be modified, but the male side needs to be cut off ready for the Draeger male 8-pin connector to be fitted.

Draeger male 8-pin side:

                     Fig 2.2  
[image: image12.wmf]Rear Housing

Connector Pin Rear View

Collett

Connector Pin

Housing

Metal Sleeving

All Keyways facing top

1

2

3

4

5

6

7

8

(1) White

(3) Blue

(5) Red

(4) Yellow

(2) Orange

(8) Black

Strain Relief

Combined 

outer and 

inner 

shields 

(is+ os) 

(Tie to collett)


1.
Feed (6 x 43mm (clear) heat shrink, strain relief, rear nut, collett and (6 x 25mm (black) heat shrink over end of cable.

2.
Strip 20mm off outer jacket of cable to reveal coloured wires, outer shield, and nylon/paper wire packing.

3.
Cut all packing and blue wire to the base.

4.
Strip 20mm off inner jacket to reveal black and white wires and the inner shield.

5.
Twist outer and inner shields together.

6.
Trim ends of wires to 13mm long.

7.
Strip jacket of every wire 2mm to reveal copper core.

8
Heat (6 x 25mm (black) heat shrink up to the end of the cable.

9.
Fold shields back along the surface and glue down over the heatshrink.

10.
Solder ends of every wire and shields to the correct positions on the rear of the connector.

11.
Slide collett over shields and to the end of the cable.

12.
Fit metal sleeves over the connector, and insert into the housing ensuring all keyways lie up.

13.
Holding the housing fixed, tighten the rear nut.

14.
Push the strain relief onto the rear nut.

TESTING 

1.
Use multi-meter to check continuity from pin to pin along the cable.

2.
Check the cable is of correct quality standard. (See VM/COP/30.11 for details).

Labelling

1.
Labels:
 to be attached facing upwards as looking at the top of the probe.


· 1 x CE Label

· 1 x Serial no. Label

· 1x Orange ‘Do Not Throw Away’ Label (correct one of two is dependant of country unit is being sold to).

[image: image13.wmf]Orange dumbell 

labell

Do Not

Throw

Away

0086

CE

Ø6 x 43 clear heat shrink

Serial No.


Quality Assurance (QA)

1.
Use multi-meter to check continuity from pin to pin along the cable.

2.
Check the cable is of correct quality standard. (See VM/COP/30.11 for details).

3.
Fill and sign attached paperwork.

Packaging 

1.
Visually check all labels are attached properly

2.
Using a twist tie (bunny clip) wrap the cable and place in a small blue Viamed plastic box, ensuring the cable is inserted in a neat and tidy presentable manor.

3.
Place a serial number sticker (supplied with the batch) on the front face of the box.

4.
Place a packed and tested sticker (also containing initials of the individual who is packing) on the right hand side top left corner of the box. Do not close box.

Final QA

1.
Final inspection. Visually ensure cable sit neatly within the box and is in a presentable state.

2.
Boxes are ready to stock in stores.
















Page 4 of 4

[image: image14.png]_1070355037.dwg

_1071041034.dwg

_1071041429.dwg

_1071042579.dwg

_1071041371.dwg

_1070355092.dwg

_1070355164.dwg

_1070354858.dwg

_1070354925.dwg

_1070354786.dwg

_1070101508.dwg

