Microstim Stimulator General Theory


Peripheral nerve stimulators have been designed to aid anesthetists detect and monitor neuromuscular function. They are used to assess the neuromuscular block during surgery and the reversal of neuromuscular block dur​ing the recovery period. By using a nerve stimulator a clinician can determine the effects of neuromuscular-blocking drugs and adjust the dosages accordingly by observing muscular re​sponse to different patterns of electrical nerve stimulation, The degree of neuromuscular relaxation is important in determining the timing of intuba​tion and extubation. 

Principles of operation

The Microstim delivers electrical cur​rent to a peripheral nerve, typi​cally the ulnar nerve at the wrist; When the peripheral nerve is stimulated suf​ficiently the muscle fibres contract. If the intensity is increased further, the stimulus is described as su​pramaximal, but muscular contraction does not in​crease. By delivering the same amount of supramaximal stimulation before and after admini​stration of neuromuscular-blocking drugs, the clini​cian can determine the effect of the neuromuscular block. 

The Microstim  supramaximal nerve stimulator is  battery  powered 

It is similar to other stimulators in offering four types of stimulation: 

Twitch, Train-of-four (TOF), Tetanic, Double burst. 

These techniques produce specific responses or pat​terns of muscle contraction. The level or type of block is assessed  by observing visually or tactilely the degree to which the affected muscle contracts. 

.

Twitch:  

Consists of a single supramaximal current and is applied repeti​tively at a frequency of 0.1 to 2 Hertz (Hz). The pulses are typically rectangular and have durations of 0.2 millisecond  or less to avoid repetitive nerve firing. Single-twitch stimulation is the least sensitive method of demonstrating a partial neuromuscular block. The twitch response is not reduced until at least 75% to 80% of the nerve endings are blocked, and it disappears completely when 90% are blocked.

The Microstim applies   Hz with a pulse duration of       milliseconds

Train of Four TOF: 

Consists of four supramaximal stimuli at 2 Hz: at half-second intervals. The amplitude of the fourth response in relation to the first gives the TOF ratio, which begins to decrease when more than 70% of the receptors are blocked. TOF stimulation can be repeated every 8 to 12 seconds.

The Microstim applies      Hz      at       intervals

Some neuromuscular blocks cannot be evaluated using Single-twitch or TOF stimulation during deep relaxation. because the associated response can disappears for extended periods e.g greater than 5 minutes

 Tetanic : 

Can consist of pulses at a much higher rate (30 to 100 Hz). 

Greater sensitivity is possible at higher frequencies. but at 100 Hz for 5 seconds a lack of appropriate response develops when 50% of the receptors are blocked; and at 200 Hz for 5 seconds  there is insufficient response when only 30% of the receptors are blocked.

Tetanic stimulation can be painful and under certain conditions may influence the course of the neuromuscular block in the investigated muscle.

The Microstim  supplies  50Hz applied for 5 seconds then 1Hz for 30 seconds

Double-burst: 

Consists of either two trains of three pulses at approximately 50 Hz, separated by approxi​mately 0.75 milliseconds,  (3.3)

or an initial train of three pulses followed by a train of two pulses (3.2).

The neuromuscular response consists of two short muscle contractions, the second of which is significantly less forceful than the first. The ratio of these contractile forces indicates the degree of patient drug-induced muscular relaxation.

The Micostim consists of 3 pulses at 50Hz followed by a 0.75 second pulse then 2 pulses at 50Hz

Application of the stimuli

Motor fibres are stimulated before pain fibres and as low stimulating currents are used for regional anesthesia the patient feels minimal discomfort
Conventional EKG disposable electrodes are most commonly used for applying the stimulation as they provide good contact with low impedance. 

Ball Electrodes

Some competitive models are also equipped with ball electrodes mounted on the stimulator  This method was rejected due to the possibility of varying skin impedance caused by the small to area of application and the conductive properties of the skin . Intermittent conductivity could cause inadvertent large jumps in the energy levels being applied.

Burns are unlikely to occur with periph​eral nerve stimulators because they use low current and typically only briefly impart electrical energy through the skin.

The Microstim was designed to be battery driven as this eliminates many of the perceived risks in relation to burns particularly with earth leakage currents and alternative paths via electrosurgical units

Microstim Specifications

Typical current        mA into 1000 ohms

Pulse width

Pulse Indicator LED

Battery type MN1604  9volt

Battery state indicator 3 colour LED

Battery capacity

Height

Width

Depth

Weight     gms          oz

Terminal polarity indication: 

Indicates the positive and negative bipolar electrodes. 

Leads comply with FDA “protected” lead set requirements

Pacemakers

Problems have been reported with use of peripheral nerve stimulators during surgical procedures. E.g pace​maker during surgery. Pacemaker function returned to normal upon cessation of peripheral nerve stimulation. Suitable warnings are given in the instructions

Associated Papers

Barthram CN. Nerve stimulators for nerve location — are they all the same? A study of stimulator per​formance. Anaesthesia 1997 Aug;52(8):761-4.

Beemer GH, Reeves JH. An evaluation of eight periph​eral nerve stimulators for monitoring neuromuscu​lar blockade. Anaesth Intensive Care 1988 Nov; 16(4):464-72.
Dankle JA, Wiegand DA. Investigation of a coaxial bipolar nerve stimulator for intraoperative motor nerve monitoring. Laryngoscope 1994 May; 104(5 Pt 1):619-22.

Ehrhardt JC, Lin CS, Magnotta VA, et al. Peripheral nerve stimulation in a whole-body echo-planar im​aging system. J Magn Reson Imaging 1997 Mar​Apr;7(2):405-9.

Gravenstein JS, Paulus DA. Monitoring practice in clinical anesthesia. 2nd ed. Philadelphia: JB Lippin​cott; 1987:178-91.

Grill WM, Mortimer JT. Non-invasive measurement of the input-output properties of peripheral nerve stimulating electrodes. J Neurosci Methods 1996 Mar;65( 1):43-50.

Hadzicá A, Vloka JD. Peripheral nerve stimulator for unassisted nerve blockade. Anesthesiology 1996 Jun;84(6):1528-9.

Myirea KC, Hameroff SR, Calkins JM, et al. Evalu​ation of peripheral nerve stimulators and relation​ship to possible errors in assessing neuromuscular blockade. Anesthesiology 1984 May;60(5):464-6.

O’Flaherty D, Wardill M, Adams AR Inadvertent sup​pression of a fixed rate ventricular pacemaker using a peripheral nerve stimulator. Anaesthesia 1993 Aug;48(8):687-9.

Rowlee SC. Monitoring neuromuscular blockade in the intensive care unit: the peripheral nerve stimulator. Heart Lung 1999 Sep-Oct;28(5):352-62.

Sansome AJ, de Courcy JG. A new dual function nerve stimulator. Anaesthesia 1989 June;44(6):494-7.

Tiel RL, Happel LT Jr, Kline DG. Nerve action poten​tial recording method and equipment. 2~eurosurgery 1996 Jul;39(1):103-8.

Weiner RL. The future of peripheral nerve neurostimu​lation. Neurol Res 2000 Apr;22(3):299-304.

Inberg P. Annila I, Annila P. Double-injection method using peripheral nerve stimulator is superior to single injection in axillary plexus block. Reg Anesth  1999 Nov-Dec;24(6):509-13.

Franco CD, Vieira ZE. 1,001 subclavian pen-vascular brachial plexus blocks: success with a nerve stimulator. Reg Anesth 2000 Jan-Feb;25( 1):41-6.

CarIes M, PulciniA, MacchiP, et al. An evalu​ation of the brachial plexus block at the humeral canal using a neurostirnulator (1417 patients): the efficacy, safety, and predictive criteria of failure. Anesth Anaig 2001 Jan;92(1):194-8.

Existing International Standards

Blue Cross and Blue Shield Association. Electrical nerve stimulation [technology assessment report]. 1988

International Electrotechnical Commission. Medical electrical equipment — part 1: general requirements for safety [standard]. IEC 60601-1 (1988-12). 1988.

Medical electrical equipment — part 1: general re​quirements for safety. Amendment 1 [standard]. IEC 60601-1-aml (1991-11). 1991.

Medical electrical equipment — part 1: general re​quirements for safety. Amendment 2 [standard].

IEC 60601-1-am2 (1995-03). 1995.

Medical electrical equipment — part 1: general re​quirements for safety. Section 2. Collateral standard:

electromagnetic compatibility — requirements and tests, IEC 60601-1-2 (1993-04). 1993.

Medical electrical equipment — part 2: particular re​quirements for the safety of nerve and muscle stimu​lators [standard]. JEC 60601-2-10 (1987-12). 1987.
The Anaestim 

This version can be used to accurately  locate peripheral nerves when performing regional anesthesia.

Needle electrodes connectect by crocodile clips transmit the electrical pulses from the stimulator to the nerve.

The correct application of electrodes is very important be​cause even slight displacements can result in consider​able changes in stimulation current requirements. In addition, the electrodes must be placed so that the device stimulates the nerve, not the muscle.

Many stimulators deliver a constant current set by the operator. Current settings typically range from 0.5 to 10 milliamperes (mA) — levels that can stimulate or depolarize a nerve without causing tissue trauma. The stimulator compensates for variations in tissue resistance and includes a digital meter to monitor the current and allow the operator to adjust the current output manually.

manager6
Page 2
22/03/2004
C:\Documents and Settings\manager6\My Documents\Stimulator Theory General.doc

