[image: image10.png]


 

                      SpO2 Extension Cable Assembly

	COMPANY OPERATING PROCEDURES


0019538  
    Odam Bruker   
         P953E8

VM3/COP/33.27
	Date:
  20-Jun-02
	Revision date: 19-Sep-03
	Issue:   3




[image: image1.wmf]Ø

6 x 43 Clear 

heatshrink. For serial 

no.

2.45m

Enlarged front view

Key-way Allignment

DPK1 Type Male 12 pin

956E8 Extension cable


[image: image2.jpg]









[image: image3.jpg]



Equipment required:
Soldering iron (0060120), solder (0050012), Wire stripper (0060030), Flush Cutter (0060010), Snipe nose pliers (0060021), ‘helping hand’ (0060145), Heat gun (0060100).

Parts list:
Kit and parts required. (Continued over page)

	SPK1-A male 12-pin Side
	D-Type Female 9-pin side

	Qty
	Description
	Part No.
	Qty
	Description
	Part No.

	
	
	
	
	
	

	1
	SPK1-A male 10-pin kit
	0010698
	1
	Pre manufactured cable
	0019568

	(1)
	
[image: image4.wmf]       Rubber housing
	kit
	
	
	

	(1)
	 
[image: image5.wmf]        Upper casing
	kit
	
	
	

	(1)
	    
[image: image6.wmf]            Pin housing
	kit
	
	
	

	1
	       
[image: image7.wmf]              Cable grip
	0010500

(100/pack)
	
	
	

	1
	(6 x 43mm Clear heat shrink
	0032331

(roll)
	
	
	

	1
	(6 x 25mm heat shrink
	0032321

(roll)
	
	
	

	1
	(1.6 x 20mm heat shrink
	0032310

(roll)
	
	
	

	
	
	
	
	
	


ASSEMBLY OPERATIONS

1.
Pre Heat soldering iron temperature to 240oc.

2.
Collect all required parts and equipment listed above.

D type 9-pin side

1.
Pre manufactured cable. Requires end of cable to be cut off to allow connector to be fitted.

SPK1-A male 12-pin Side :
Fig 2.2  
[image: image8.wmf]Rear View of Pin Housing

(12) Black

(1) Red

(4) Orange

(8) Yellow

(3) White

(7) Blue

1

Pin Housing

Casing

6

8

7

12

9

4

5

2

10

11

3

(5) 

Combined 

outer and 

inner 

shields 

(is+ os) 

Rubber Housing

Cable grip


1.
Feed (6 x 43mm (clear) heat shrink, rubber housing, casing, cable grip and (6 x 25mm (black) heat shrink over the end of the cable.

2.
Strip 25mm off outer jacket of cable to reveal coloured wires, outer shield, and nylon/paper wire packing.

3.
Cut all packing to the base.

4.
Strip 25mm off inner jacket to reveal black and white wires and the inner shield.

5.
Twist inner and outer shields together, and heat (6 x 25mm heat shrink over excess naked wire.

6.
Trim ends of wires to the same length.

7.
Strip jacket of every wire 2mm to reveal copper core, and solder wires to the correct pin locations. 

8.
Clamp cable grip at the end of the cable.

9.
Push pin housing into the casing,.

10.
Push rubber housing over casing.

TESTING 

1.
Use multi-meter to check continuity between each connector pin.
Labelling

1.
Labels: to be attached facing upwards as looking at the top of the probe.


· 1 x CE Label

· 1 x Serial no. Label

· 1x Orange ‘Do Not Throw Away’ Label (correct one of two is dependant of country unit is being sold to).

[image: image9.wmf]0086

CE

Orange dumbell 

labell

Do Not

Throw

Away

Ø6 x 43 clear heat shrink

Serial No.


Quality Assurance (QA)

1.
Use multi-meter to check continuity between each connector pin.
2.
Fill and sign attached paperwork.

Packaging 

1.
Visually check all labels are attached properly

2.
Using a twist tie (bunny clip) wrap the cable and place in a small blue Viamed plastic box, ensuring the cable is inserted in a neat and tidy presentable manor.

3.
Place a serial number sticker (supplied with the batch) on the front face of the box.

4.
Place a packed and tested sticker (also containing initials of the individual who is packing) on the right hand side top left corner of the box. Do not close box.

Final QA

1.
Final inspection. Visually ensure cable sit neatly within the box and is in a presentable state.

2.
Boxes are ready to stock in stores.
















Page 3 of 3

[image: image10.png]_1086159163.dwg

_1091444873.dwg

_1086160197.dwg

_1086157821.dwg

_1086159129.dwg

_1069494359.dwg

_1086093342.dwg

