[image: image24.png]


 

                      SpO2 Extension Cable Assembly

	COMPANY OPERATING PROCEDURES


0019260
SPACELABS
P926E10

VM3/COP/33.02
	Date:
  13-Dec-01
	Revision date: 19-Sep-03
	Issue:   3




[image: image1.wmf]Hypertronics Male 12 Pin

D-Type Female 9 Pin

Ø

6 x 43 Clear 

heatshrink. For serial 

no.

3m

Enlarged front view


[image: image2.jpg]


                                                          [image: image3.jpg]



Equipment required:
Soldering iron (0060120), solder (0050012), Wire stripper (0060030), Flush Cutter (0060010), Snipe nose pliers (0060021), ‘helping hand’ (0060145), Heat gun (0060100).

Parts list:
Kit and parts required. (Continued over page)

	D-Type female 

9-pin Side
	Hypertronics male 

12-pin Side

	Qty
	Description
	Part No.
	Qty
	Description
	Part No.

	
	
	
	
	
	

	1
	D-type female 9-pin kit
	0010760
	1
	Hypertronics male 

12-pin Kit
	0010602

	(1)
	
[image: image4.wmf]  Outer Casing
	kit
	(1)
	  
[image: image5.wmf]         Pin Housing
	kit

	(1)
	       
[image: image6.wmf]              Cable grip
	kit
	(12)
	
[image: image7.wmf]           Pins
	kit

	(1)
	
[image: image8.wmf]           Pin Housing
	kit
	(1)
	       
[image: image9.wmf]              Cable grip
	kit

	(9)
	
[image: image10.wmf]           Pins
	kit
	(1)
	  
[image: image11.wmf]       Strain relief
	kit

	1
	1.3m 6-core cable
	0030513

(roll)
	(1)
	   
[image: image12.wmf]            Collett
	kit

	1
	(1.6 x 17mm heat shrink
	0032310

(roll)
	(1)
	  
[image: image13.wmf]           Barrel
	kit

	1
	(6 x 10mm heat shrink
	0032321

(roll)
	1
	
[image: image14.wmf]R1

     10 ( Resistor
	0032010

	
	
	
	1
	
[image: image15.wmf]R2

  60.4 k( Resistor
	0032130

	
	
	
	1
	(6 x 43mm Clear heat shrink
	0032331

(roll)

	
	
	
	1
	(6 x 10mm heat shrink
	0032321

(roll)

	
	
	
	1
	(1.6 x 4mm heat shrink
	0032310

(roll)


ASSEMBLY OPERATIONS

1.
Pre Heat soldering iron temperature to 240oc.

2.
Collect all required parts and equipment listed above.

3.
Cut a 3.1 metre length of standard 6-core cable. (Details shown below).

                           Fig 1.  
[image: image16.wmf]1.2 Cable Cross Section

1.0 Mechanical Specifications

1.1 Cable Schematic

Black

White

Blue

Inner Shield

Orange

Yellow

Red

Outer Shield

Outer Jacket

Outer Shield

Red,Orange, Yellow, Blue

Inner Shield

Black, White Twisted Pair

Twisted Pair

Paper/ nylon packing to make round


D-Type female 9-pin side:

                   Fig 2.1 
[image: image17.wmf]Pin Insert Rear View

1

Connector Plug

(2) Red

(3) Yellow

(1) Blue

(6) Orange

Pins

(9) Black

(5) White

Cable Grip

Outer Casing

6

7

8

9

4

5

2

3

(7) 

Combined 

outer and 

inner 

shields 

(is+ os) 


1.
Feed outer casing, cable grip and (6 x 10mm heat shrink (black) over the end of the cable.

2.
Strip 20mm off outer jacket of cable to reveal coloured wires, outer shield, and nylon/paper  packing.

3.
Cut all packing to the base.

4.
Strip 20mm off inner jacket to reveal black and white wires and the inner shield.

5.
Twist outer and inner shields together.

6.
Trim ends of wires and shields to the same length.

7.
Strip jacket of every wire 2mm to reveal copper core.

8.
Heat (1.6 x 17mm heat shrink over twisted inner and outer shields to cover naked wire, and solder end to the rear of one pin.

9.
Apply small amount of solder to ends of each wire and solder each of the 6 wires to the rear separate pins.

10.
Clamp cable grip approximately 2mm from end of outer jacket.

11.
Place (6 x 10mm heat shrink over cable grip and beginning of wires and heat to shrink firmly over.

12.
Insert pins into correct locations (as shown in fig 2.1) and push firmly into place.

13.
Push outer casing over cable grip and wires to fit around the pin housing.

Hypertronics male 12 pin side:

Fig 2.2  
[image: image18.wmf](2) Black

(1) White

(3) Orange

(11) Yellow

(8) Blue

(12) Red

Pins

1

2

3

4

5

6

7

9

8

10

11

12

Pin Housing Rear View

Collett

End Cap

Strain Relief

Cable Grip

Pin Housing

Barrel

R1

W

(3,2) R1 = 10

R2

(3,7) R2 = 60.4

K

W

(3) 

Combined 

outer and 

inner 

shields 

(is+ os) 


1.
Feed (6 x 43mm (clear) heat shrink, end cap, collett, strain relief, (6 x 10mm heat shrink (black) and cable grip over end of wire.

2.
Strip 20mm off outer jacket of wire- to reveal coloured wires, outer shield, and nylon/paper wire packing.

3.
Cut all packing to the base.

4.
Strip 20mm off inner jacket of wire- to reveal black and white wires and the inner shield.

5.
Twist outer and inner shields together and trim with orange wire to approximately 6mm long.

6.
Trim (tidy) all ends of remaining wires to the same length.

7.
Strip jacket of every wire 2mm to reveal copper core.

8.
Apply small amount of solder to the ends of each wire and shields.

9.
Solder orange wire to the end of the twisted shield pair.

10.
Heat (1.6 x 4mm heat shrink over twisted shield pair and orange wire to insulate.

11.
Cut one leg of the (R1) 10( resistor to 4mm and the other to 15mm, and solder each side to the rear of 2 separate pins.

12.
Solder the black wire to the pin with short 4mm side of the 10( resistor.

13.
Cut one leg of (R2) 60.4 k( resistor to 4mm and solder to the rear on one pin.

14.
Solder remaining wires to the rear of pins.

15.
Push/pull all pins firmly into the correct locations shown in fig 2.2.

16.
Solder remaining leg of (R2) to the long leg of (R1) as shown in fig 2.2. and trim off excess wire. (form wire to shape, solder then trim to be as compact as possible). 

17.
Clamp cable grip approximately 2mm from outer jacket end.

18.
Place (6 x 10mm heat shrink over cable grip and beginning of wires and heat to shrink firmly around.

19.
Push strain relief up to cable grip, collett over strain relief and up to pin housing and into the barrel and finally screw end cap onto the barrel.

TESTING 

1.
Attach male 12-pin side to a Spacelabs/Nellcor adapter cable and then to the test box connector labelled (A).

2.
Attach female 7-pin side to a Novametrics probe.

3.
Check display is showing correct characteristic as shown below. (At correct switch positions)

[image: image19.jpg]



[image: image20.jpg]R -

—— TN m—— T~

e

A AR e




                              NB: Gap to be at bottom of display


            
            Position 2. IR, LED.
       
      Position 4. Detector

4.
If gap is at the top of the screen then LED is wired the wrong way around.

5.
‘Play’ with wire at connections to see if any change in the display (flickering etc).

6.
If there is any movement of signal, the extension wire must be taken apart and all connections checked and re-soldered. Then tested again, until results are satisfactory.

7.
Check the cable is of correct quality standard. (See VM/COP/30.11 for details).

8.
Connect cable to the test cable and then the Nellcor monitor, and attach probe to the finger, check SpO2 level (ideal reading between 95-100).

Labelling

1.
Labels: to be attached facing upwards as looking at the top of the probe.


· 1 x CE Label

· 1 x serial no. Label

· 1x Orange ‘Do Not Throw Away’ Label (correct one of two is dependant of country unit is being sold to).

[image: image21.wmf]Serial No.

Ø6 x 43 clear heat shrink

Do Not

Throw

Away

CE

0086

Orange dumbell 

labell


Quality Assurance (QA)

1.
Attach Hypertronics male 12-pin side to a Spacelabs/Nellcor adapter cable and then to the test box connector labelled (A).

2.
Attach D-type female 9-pin side to a Novametrics probe.

3.
Check display is showing correct characteristic as shown below. (At correct switch positions)

[image: image22.jpg]



[image: image23.jpg]R -

—— TN m—— T~

e

A AR e




                              NB: Gap to be at bottom of display


            
            Position 2. IR, LED.
       
      Position 4. Detector

4.
If gap is at the top of the screen then LED is wired the wrong way around.

5.
‘Play’ with wire at connections to see if any change in the display (flickering etc).

6.
If there is any movement of signal, the extension wire must be taken apart and all connections checked and re-soldered. Then tested again, until results are satisfactory.

7.
Check the cable is of correct quality standard. (See VM/COP/30.11 for details).

8.
Connect cable to the test cable and then the Nellcor monitor, and attach probe to the finger, check SpO2 level (ideal reading between 95-100).

9.
Fill and sign attached paperwork.

Packaging 

1.
Visually check all labels are attached properly

2.
Using a twist tie (bunny clip) wrap the cable and place in a small blue Viamed plastic box, ensuring the cable is inserted in a neat and tidy presentable manor.

3.
Place a serial number sticker (supplied with the batch) on the front face of the box.

4.
Place a packed and tested sticker (also containing initials of the individual who is packing) on the right hand side top left corner of the box. Do not close box.

Final QA

1.
Final inspection. Visually ensure cable sit neatly within the box and is in a presentable state.

2.
Boxes are ready to stock in stores.
















Page 6 of 6

[image: image24.png]_1069495046.dwg

_1070093120.dwg

_1073376711.dwg

_1075720021.dwg

_1075719936.dwg

_1070093428.dwg

_1069504476.dwg

_1069763934.dwg

_1069766336.dwg

_1069744209.dwg

_1069495127.dwg

_1069494359.dwg

_1069494489.dwg

_1069494795.dwg

_1069494204.dwg

