[image: image19.png]


 

                      SpO2 Extension Cable Assembly

	COMPANY OPERATING PROCEDURES


0019278
SPACELABS
P927E8

VM3/COP/33.03
	Date:
  13-Dec-01
	Revision date: 8-Dec-03
	Issue:   3




[image: image1.wmf]Enlarged front view

P956E8 Extension cable

Ø

6 x 43 Clear 

heatshrink. For serial 

no.

2.4m

Hypertronics Male 12 Pin


[image: image2.jpg]


                                                          [image: image3.jpg]



Equipment required:
Soldering iron (0060120), solder (0050012), Wire stripper (0060030), Flush Cutter (0060010), Snipe nose pliers (0060021), ‘helping hand’ (0060145), Heat gun (0060100).

Parts list:
Kit and parts required. (Continued over page)

	D-Type female 

9-pin Side
	Hypertronics male 

12-pin Side

	Qty
	Description
	Part No.
	Qty
	Description
	Part No.

	
	
	
	
	
	

	1
	D-type extension cable

(Female side)
	0019568
	1
	Kit
	0010602

	
	
	
	(1)
	  
[image: image4.wmf]         Pin Housing
	kit

	
	
	
	(12)
	
[image: image5.wmf]           Pins
	kit

	
	
	
	(1)
	       
[image: image6.wmf]              Cable grip
	kit

	
	
	
	(1)
	  
[image: image7.wmf]       Strain relief
	kit

	
	
	
	(1)
	   
[image: image8.wmf]            Collett
	kit

	
	
	
	(1)
	  
[image: image9.wmf]           Barrel
	kit

	
	
	
	1
	
[image: image10.wmf]R1

     10 ( Resistor
	0032010

	
	
	
	1
	
[image: image11.wmf]R2

  60.4 k( Resistor
	0032130

	
	
	
	1
	(6 x 43mm Clear heat shrink
	0032331

(roll)

	
	
	
	1
	(6 x 10mm heat shrink
	0032321

(roll)

	
	
	
	1
	(1.6 x 4mm heat shrink
	0032310

(roll)


ASSEMBLY OPERATIONS

1.
Pre Heat soldering iron temperature to 240oc.

2.
Collect all required parts and equipment listed above.

D-type female 9-pin side:

  
[image: image12.wmf]2.4m

D-Type Female 9 Pin

D-Type Male 9 Pin


 Fig 2.1.  0019568 Extension Cable.
1.
The 0019568 cable is already assembled. The female side does not need to be modified, but the male side needs to be cut off ready for the Hypertronics male 12-pin connector to be fitted.


Hypertronics male 12 pin side:

Fig 2.2  
[image: image13.wmf](3) 

Combined 

outer and 

inner 

shields 

(is+ os) 

Cable Grip

End Cap

Collett

Strain Relief

(11) Red

(8) Blue

(12) Yellow

(3) Orange

Pins

(1) White

(2) Black

Pin Housing

Barrel

(3,7) R2 = 60.4

(3,2) R1 = 10

W

K

W

R2

R1

9

2

5

4

11

10

3

7

6

12

8

1

Pin Housing Rear View


1.
Feed (6 x 43mm (clear) heat shrink, end cap, collett, strain relief, (6 x 10mm heat shrink (black) and cable grip over end of wire.

2.
Strip 20mm off outer jacket of wire- to reveal coloured wires, outer shield, and nylon/paper wire packing.

3.
Cut all packing, and green (if there is one) wire to the base.

4.
Strip 20mm off inner jacket of wire- to reveal black and white wires and the inner shield.

5.
Twist outer and inner shields together and trim with orange wire to approximately 6mm long.

6.
Trim (tidy) all ends of remaining wires to the same length.

7.
Strip jacket of every wire 2mm to reveal copper core.

8.
Apply small amount of solder to the ends of each wire and shields.

9.
Solder orange wire to the end of the twisted shield pair.

10.
Heat (1.6 x 4mm heat shrink over twisted shield pair and orange wire to insulate.

11.
Cut one leg of the (R1) 10(  resistor to 4mm and the other to 15mm, and solder each side to the rear of 2 separate pins.

12.
Solder the black wire to the pin on the short 4mm side of the 10(  resistor.

13.
Cut one leg of (R2) 60.4 k( resistor to 4mm and solder to the rear on one pin.

14.
Solder remaining wires to the rear of pins.

15.
Push/pull all pins firmly into the correct locations shown in fig 2.2.

16.
Solder remaining leg of (R2) to the long leg of (R1) as shown in fig 2.2. and trim off excess wire. (form wire to shape, solder then trim to be as compact as possible). 

17.
Solder the orange and twisted pair of shields to R1 pin 3. 

18.
Clamp cable grip approximately 2mm from outer jacket end.

19.
Place (6 x 10mm heat shrink over cable grip and beginning of wires and heat to shrink firmly around.

20.
Push strain relief up to cable grip, collett over strain relief and up to pin housing and into the barrel and finally screw end cap onto the barrel.

TESTING

1.
Attach male 12-pin side to a Spacelabs/Nellcor adapter cable and then to the test box connector labelled (A).

2.
Attach female 9-pin side to a Nellcor probe.

3.
Check display is showing correct characteristic as shown below. (At correct switch positions)

[image: image14.jpg]



[image: image15.jpg]R -

—— TN m—— T~

e

A AR e




                              NB: Gap to be at bottom of display


            
            Position 2. IR, LED.
       
      Position 4. Detector

4.
If gap is at the top of the screen then LED is wired the wrong way around.

5.
‘Play’ with wire at connections to see if any change in the display (flickering etc).

6.
If there is any movement of signal, the extension wire must be taken apart and all connections checked and re-soldered. Then tested again, until results are satisfactory.

7.
Check the cable is of correct quality standard. (See VM/COP/30.11 for details).

8.
Connect cable to a test lead then the Nellcor monitor, and attach probe to the finger, check SpO2 level (ideal reading between 95-100).

Labelling

1.
Labels: to be attached facing upwards as looking at the top of the probe.


· 1 x CE Label

· 1 x serial no. Label

· 1x Orange ‘Do Not Throw Away’ Label (correct one of two is dependant of country unit is being sold to).

[image: image16.wmf]Serial No.

Ø6 x 43 clear heat shrink

Do Not

Throw

Away

CE

0086

Orange dumbell 

labell


Quality Assurance (QA)

1.
Attach male 12-pin side to a Spacelabs/Nellcor adapter cable and then to the test box connector labelled (A).

2.
Attach female 9-pin side to a Nellcor probe.

3.
Check display is showing correct characteristic as shown below. (At correct switch positions)

[image: image17.jpg]



[image: image18.jpg]R -

—— TN m—— T~

e

A AR e




                              NB: Gap to be at bottom of display


            
            Position 2. IR, LED.
       
      Position 4. Detector

4.
If gap is at the top of the screen then LED is wired the wrong way around.

5.
‘Play’ with wire at connections to see if any change in the display (flickering etc).

6.
If there is any movement of signal, the extension wire must be taken apart and all connections checked and re-soldered. Then tested again, until results are satisfactory.

7.
Check the cable is of correct quality standard. (See VM/COP/30.11 for details).

8.
Connect cable to a test lead then the Nellcor monitor, and attach probe to the finger, check SpO2 level (ideal reading between 95-100).

9.
Fill and sign attached paperwork.

Packaging 

1.
Visually check all labels are attached properly

2.
Using a twist tie (bunny clip) wrap the cable and place in a small blue Viamed plastic box, ensuring the cable is inserted in a neat and tidy presentable manor.

3.
Place a serial number sticker (supplied with the batch) on the front face of the box.

4.
Place a packed and tested sticker (also containing initials of the individual who is packing) on the right hand side top left corner of the box. Do not close box.

Final QA

1.
Final inspection. Visually ensure cable sit neatly within the box and is in a presentable state.

2.
Boxes are ready to stock in stores.
















Page 3 of 5

[image: image19.png]_1069495046.dwg

_1069504476.dwg

_1069848476.dwg

_1069851405.dwg

_1132390017.dwg

_1069766336.dwg

_1069763934.dwg

_1069495127.dwg

_1069494359.dwg

_1069494489.dwg

_1069494795.dwg

_1069494204.dwg

