Certificate of Calibration.

Certificate .
Date of Calibration: 


Manufacturer: 



Viamed Ltd.


Description: 



DL-3000 SpO2 Simulator.


Serial No: 



.


Software Version: 


.



PCB Issue: 



.

The DL-3000 was calibrated and the following checked:

1
Internal connections.








  

2
LCD functioning. 








3
Fascia LEDs functioning. 









4
User selectors functioning.









5
SpO2 simulation.









6
BPM simulation.








7
ECG simulation.









8
Electrical and mechanical safety.











SpO2 and BPM simulation results as checked on a Nellcor NBP-40 oximeter:

DL-3000 Settings

SpO2
BPM

100
97
90
80
200
180
160
140
120
100
80
60
40















100
97
90
79
199
180
160
140
120
100
80
60
40

Functional test & calibration carried out by:



(Peter Anderson)

This is to certify that the equipment detailed hereon has been calibrated and now conforms to Viamed Limited published specifications and where applicable BS EN 61010.  The Q.A. arrangements for the above are in strict conformity to ISO 9001/2000.  All measurements made are traceable to NAMAS.

