[image: image3.jpg]


Instructions For Use

[image: image4.jpg]



NeoPEEP®
 
Single Patient Use 
Neonatal Resuscitation Circuit

· Ensure the NeoPEEP® neonatal resuscitation circuit is complete and undamaged.
· Set up the resuscitation unit as per the manufacturer’s guidelines, and connect NeoPEEP® Circuit to patient outlet port (A).

· Adjust gas supply (oxygen or blended oxygen/air) to required flow rate as defined by the hospital’s protocol for resuscitation.

· Occlude the patient opening of the ‘T-Piece’ (B) and the PEEP valve outlet (C) to create an air tight seal. 

· Adjust the PIP* pressure on the resuscitation device to the required pressure.

· Uncover the PEEP valve outlet (C). Adjust PEEP* setting by adjusting the PEEP control cap (D) on the NeoPEEP® Circuit until the reading on the resuscitation device pressure gauge indicates that the correct PEEP, as defined by the hospital’s protocol for resuscitation, has been achieved.


*Note: 
PIP = Peak Inspiratory Pressure   

PEEP = Positive End Expiratory Pressure
[image: image1.png]


  Resuscitation
· If required, connect the NeoPEEP® Circuit to appropriate size ET tube or Laryngeal Mask. If a mask is to be used, connect the NeoPEEP® Circuit to a neonatal resuscitation mask of an appropriate size for the patient.

· Place the mask over the patient’s airway as per resuscitation training guidelines, ensuring a good seal around the airway.

· Resuscitate by occluding the PEEP valve outlet (C) with thumb to allow inspiration. 

· Uncover the PEEP valve outlet (C) to allow expiration. 

· If necessary adjust the PEEP setting by turning the PEEP control cap (D) accordingly 

· Repeat steps for inspiration and expiration as necessary, following the hospital protocol for resuscitation.

NeoPEEP® Nominal PEEP Values

5 l/min
= 0.1 - 10cm H2O/mbar


Operating Environmental Temperature Limits: -18ºC to +50ºC
10 l/min
= 0.7 - 35cm H2O/mbar


Storage Environmental Temperature Limits:     -40ºC to +60ºC

15 l/min
= 1.5 - 45cm H2O/mbar

[image: image5.jpg]0434



[image: image6.jpg]szrsh all

Excellence as Standard




	


[image: image7.png]


[image: image8.jpg]0434



[image: image2.png]



CAUTION: This circuit must only be used by competent, trained personnel in accordance with local resuscitation guidelines.

�





� INCLUDEPICTURE "http://www.abbottdiagnostics.com/Customer_Support/Customer_Communications/img/cc_noreuse.gif" \* MERGEFORMATINET ���





C





A





B





D





�








	Manufactured by:
	
	Distributed by:
	
	
	

	[image: image3.jpg]
	Marshall Products Ltd.

1 The Maltings, Brassmill Lane, Bath, 

BA1 3JL, United Kingdom

Tel: 0845 612 8888 - Fax: 0845 612 8889

E-mail: sales@marshallproducts.co.uk

www.marshallproducts.co.uk
	[image: image4.jpg]
	Viamed

15 Station Road, Cross Hills, Keighley, 

West Yorkshire BD20 7DT United Kingdom

Tel: +44(0)1535 634542 Fax: +44(0)1535 635582

Email: info@viamed.co.uk

www.viamed.co.uk
	
	[image: image5.jpg]


Rev: 02/08/07


