[image: image9.png]


 

                           SpO2 ‘Y’ Probe Assembly

	COMPANY OPERATING PROCEDURES


0018612    
   BCI   
         P861YA

VM3/COP/35.08
	Date:
  14-May-02
	Revision date: 19-Sep-03
	Issue:   3




[image: image1.wmf]0.9m

D-Type Male 9 Pin

Y- Probe

TOP


[image: image2.jpg]





             

[image: image3.jpg]



Equipment required:
Soldering iron (0060120), solder (0050012), Wire stripper (0060030), Flush Cutter (0060010), Snipe nose pliers (0060021), ‘helping hand’ (0060145), Heat gun (0060100).

Parts list:
Kit and parts required. (Continued over page)

	D-type male 9-pin

Pre-manufactured
	Y Probe Side

Pre-manufactured

	Qty
	
	Part No.
	Qty
	Description
	Part No.

	
	
	
	
	
	

	1
	Pre- manufactured probes.
	0018612

	
	
	
	
	
	

	
	
	
	
	
	


ASSEMBLY OPERATIONS

1.
P861YA is a pre-manufactured unit that is not assembled by Viamed.

TESTING 

1.
Attach D-Type male 9-pin side to the test box connector marked  ‘A’.

2.
Check display is showing correct characteristics as shown below. (At correct switch positions)

[image: image9.png]
      LED should read approx 1.8v

[image: image4.jpg]


                    
[image: image5.jpg]R -

—— TN m—— T~

e

A AR e



                     


            
     Position 1. IR, LED.
       
                           Position 4. Detector

3.
If the LED signal is at the bottom then it is wired incorrectly.

4.
‘Play’ with wire at connections to see if any change in the display (i.e. flickering etc).

5.
If there is any movement of signal, the cable is faulty it needs to be reported to the senior worker. Then tested again until results are satisfactory.

6.
Check the cable is of correct quality standard. (See VM/COP/30.11 for details).

7.
Attach D-Type male 9-pin side to a 71000A2 monitor and the probe on to the ear to check SpO2 level. (Ideal reading 95-100.)

Labelling

1.
Labels: to be attached facing upwards as looking at the top of the probe.


· 1 x CE Label

· 1x Orange / Brown BCI Label 


[image: image6.wmf]For use on

BCI

Compatible

Pulse Oximeters

Orange/ Brown 

dumbell labell

0086

CE


Quality Assurance (QA)

1.
Attach D-Type male 9-pin side to the test box connector marked  ‘A’.

2.
Check display is showing correct characteristics as shown below. (At correct switch positions)


     LED should read approx 1.8v

[image: image7.jpg]


                    
[image: image8.jpg]R -

—— TN m—— T~

e

A AR e





            
     Position 1. IR, LED.
       
                           Position 4. Detector

3.
If the LED signal is at the bottom then it is wired incorrectly.

4.
‘Play’ with wire at connections to see if any change in the display (i.e. flickering etc).

5.
If there is any movement of signal, the cable is faulty it needs to be reported to the senior worker. Then tested again until results are satisfactory.

6.
Check the cable is of correct quality standard. (See VM/COP/30.11 for details).

7.
Attach D-Type male 9-pin side to a 71000A2 monitor and the probe on to the ear to check SpO2 level. (Ideal reading 95-100.)

8.
Fill and sign attached paperwork.

Packaging 

1.
Visually check all labels are attached properly

2.
Using a twist tie (bunny clip) wrap the cable and place in a small blue Viamed plastic box, ensuring the cable is inserted in a neat and tidy presentable manor.

3.
Place a serial number sticker (supplied with the batch) on the front face of the box.

4.
Place a packed and tested sticker (also containing initials of the individual who is packing) on the right hand side top left corner of the box. Do not close box.

Final QA

1.
Final inspection. Visually ensure cable sit neatly within the box and is in a presentable state.

2.
Boxes are ready to stock in stores.
















Page 2 of 3

_1082889607.dwg

_1083573038.dwg

