[image: image17.png]


 

                                SpO2 ‘Y’ Probe Assembly

	COMPANY OPERATING PROCEDURES


0018662   
Simed Baxter    
P866YA

VM3/COP/35.13

	Date:
  14-May-02
	Revision date: 19-Sep-03
	Issue:   3




[image: image1.wmf]TOP

Y- Probe

3.65m

Hypertronics Male 9 Pin

Enlarged front view


[image: image2.jpg]




    
                                [image: image3.jpg]



Equipment required:
Soldering iron (0060120), solder (0050012), Wire stripper (0060030), Flush Cutter (0060010), Snipe nose pliers (0060021), ‘helping hand’ (0060145), Heat gun (0060100).

Parts list:
Kit and parts required. (Continued over page)

	Hypertronics male 9-pin Kit
	‘Y’ Probe Side

	Qty
	
	Part No.
	Qty
	Description
	Part No.

	
	
	
	
	
	

	1
	Hypertronics male 

9-pin Kit
	0010601
	1
	Pre manufactured ‘Y’probe cable
	0018582

	(1)
	  
[image: image4.wmf]         Pin Housing
	kit
	
	
	

	(12)
	
[image: image5.wmf]           Pins
	kit
	
	
	

	(1)
	       
[image: image6.wmf]              Cable grip
	kit
	
	
	

	(1)
	  
[image: image7.wmf]       Strain relief
	kit
	
	
	

	(1)
	   
[image: image8.wmf]            Collett
	kit
	
	
	

	(1)
	  
[image: image9.wmf]           Barrel
	kit
	
	
	

	1
	
[image: image10.wmf]R1

     1.58 k( Resistor
	0032050
	
	
	

	1
	(6 x 43mm Clear heat shrink
	0032331
	
	
	

	1
	(6 x 10mm heat shrink
	0032321
	
	
	


ASSEMBLY OPERATIONS

1.
Pre Heat soldering iron temperature to 240oc.

2.
Collect all required parts and equipment listed above.

‘Y’ Probe side:

1.
Probe side is pre-manufactured and ready to have the connector assembled to it.

Hypertronics male 9 pin side:

Fig 2.2  
[image: image11.wmf](3 & 4) R1 = 1.58 k

Barrell

R1

6

4

5

Pin Housing

(-) Orange

(3) Yellow

(-) Blue

Pins

(2) Red

(6) Black

(7) White

Pin Housing Rear View

9

1

2

3

8

7

(1 & 9) 

Combined 

outer and 

inner 

shields 

(is+ os) 

Cable grip

Strain Relief

End Cap

Collett

W


1.
Cut off the d-type connector from the end of the cable.

2.
Feed (6 x 43mm (clear) heat shrink, end cap, collett, strain relief, (6 x 10mm (black) heat shrink and cable grip onto the cable.

3.
Strip 20mm off outer jacket of wire- to reveal 4 coloured wires, outer shield, and nylon/paper wire packing.

4.
Cut all packing, orange and blue wire to the base.

5.
Strip 20mm off inner jacket to reveal black and white wires and the inner shield.

6.
Twist outer and inner shields together and trim to approximately 8mm long.

7.
Strip jacket of every wire 2mm to reveal copper core.

8.
Apply small amount of solder to ends of each wire and shields.

9.
Trim one of the legs of the resistor to 4mm and the other to 15mm .

10.
Solder each leg of the resistor into the rear of 2 separate pins and push pins firmly into correct locations as shown in   fig 2.2.

11.
Solder the red wire to the short side of the resistor, which is to go into location 3, and remaining 3 wires and shields to the rear of 4 pins and push/pull firmly into correct locations shown in fig 2.2.

12.
Clamp cable grip approximately 2mm from outer jacket end.

13.
Place (6 x 10mm heat shrink over cable grip and beginning of wires and heat to shrink firmly around.

14.
Push the strain relief up to the cable grip, collett over the strain relief up to the pin housing, and into the barrel and finally screw end cap onto the barrel.

TESTING 

1.
Attach Hypertronics male 9 pin side to the test box connector marked ‘L’.

2.
Check display is showing correct characteristics as shown below. (At correct switch positions)

[image: image17.png]
      LED should read approx 1.8v

[image: image12.jpg]


                    
[image: image13.jpg]R -

—— TN m—— T~

e

A AR e





            
     Position 1. IR, LED.
       
                           Position 4. Detector

3.
If the LED signal is at the bottom then it is wired incorrectly.

4.
‘Play’ with wire at connections to see if any change in the display (i.e. flickering etc).

5.
If there is any movement of signal, the cable must be taken apart and all connections checked and re-soldered. Then tested again until results are satisfactory.

6.
Check the cable is of correct quality standard. (See VM/COP/30.11 for details).

7.
Attach Hypertronics male 9 pin side to an S100E monitor and the probe on to the ear to check SpO2 level. (Ideal reading 95-100.)

Labelling

1.
Labels:
 to be attached facing upwards as looking at the top of the probe.


· 1 x CE Label

· 1x Orange ‘Do Not Throw Away’ Label (correct one of two is dependant of country unit is being sold to).

[image: image14.wmf]Do Not

Throw

Away

Orange dumbell 

labell

CE

0086


Quality Assurance (QA)

1.
Attach Hypertronics male 9-pin side to the test box connector marked ‘L’.

2.
Check display is showing correct characteristics as shown below. (At correct switch positions)


      LED should read approx 1.8v

[image: image15.jpg]


                    
[image: image16.jpg]R -

—— TN m—— T~

e

A AR e





            
     Position 1. IR, LED.
       
                           Position 4. Detector

3.
If the LED signal is at the bottom then it is wired incorrectly.

4.
‘Play’ with wire at connections to see if any change in the display (i.e. flickering etc).

5.
If there is any movement of signal, the cable must be taken apart and all connections checked and re-soldered. Then tested again until results are satisfactory.

6.
Check the cable is of correct quality standard. (See VM/COP/30.11 for details).

7.
Attach Hypertronics male 9-pin side to an S100E monitor and the probe on to the ear to check SpO2 level. (Ideal reading 95-100.)

8.
Fill and sign attached paperwork.

Packaging 

1.
Visually check all labels are attached properly

2.
Using a twist tie (bunny clip) wrap the cable and place in a small blue Viamed plastic box, ensuring the cable is inserted in a neat and tidy presentable manor.

3.
Place a serial number sticker (supplied with the batch) on the front face of the box.

4.
Place a packed and tested sticker (also containing initials of the individual who is packing) on the right hand side top left corner of the box. Do not close box.

Final QA

1.
Final inspection. Visually ensure cable sit neatly within the box and is in a presentable state.

2.
Boxes are ready to stock in stores.
















Page 2 of 4

_1069495046.dwg

_1069504476.dwg

_1083502292.dwg

_1083502414.dwg

_1082896171.dwg

_1069495127.dwg

_1069494359.dwg

_1069494489.dwg

_1069494795.dwg

_1069494204.dwg

