[image: image16.png]


 

                                SpO2 ‘Y’ Probe Assembly

	COMPANY OPERATING PROCEDURES


0018672   
Ohmeda    
P867YA

VM3/COP/35.14

	Date:
  14-May-02
	Revision date: 10-Sep-04
	Issue:   3




[image: image1.wmf]TOP

Y- Probe

3.65m

Hypertronics Male 9 Pin

Enlarged front view


[image: image2.jpg]




    
                                           [image: image3.jpg]



Equipment required:
Soldering iron (0060120), solder (0050012), Wire stripper (0060030), Flush Cutter (0060010), Snipe nose pliers (0060021), ‘helping hand’ (0060145), Heat gun (0060100).

Parts list:
Kit and parts required. (Continued over page)

	Hypertronics male 9-pin Kit
	‘Y’ Probe Side

	Qty
	
	Part No.
	Qty
	Description
	Part No.

	
	
	
	
	
	

	1
	Hypertronics male 

9-pin Kit
	0010604
	1
	Pre manufactured ‘Y’probe cable
	0018672

	(1)
	  
[image: image4.wmf]         Pin Housing
	kit
	
	
	

	(12)
	
[image: image5.wmf]           Pins
	kit
	
	
	

	(1)
	       
[image: image6.wmf]              Cable grip
	kit
	
	
	

	(1)
	  
[image: image7.wmf]       Strain relief
	kit
	
	
	

	(1)
	   
[image: image8.wmf]            Collett
	kit
	
	
	

	(1)
	  
[image: image9.wmf]           Barrel
	kit
	
	
	

	1
	
[image: image10.wmf]R1

     1.58 k( Resistor
	0032050
	
	
	

	1
	(6 x 43mm Clear heat shrink
	0032331
	
	
	

	1
	(6 x 10mm heat shrink
	0032321
	
	
	


ASSEMBLY OPERATIONS

1.
P867YA is a pre-manufactured unit that is not assembled by Viamed.

TESTING 

1.
Attach Hypertronics male 9 pin side to the test box connector marked ‘B’.

2.
Check display is showing correct characteristics as shown below. (At correct switch positions)

[image: image16.png]
      LED should read approx 1.8v

[image: image11.jpg]


                    
[image: image12.jpg]R -

—— TN m—— T~

e

A AR e





            
     Position 1. IR, LED.
       
                           Position 4. Detector

3.
If the LED signal is at the bottom then it is wired incorrectly.

4.
‘Play’ with wire at connections to see if any change in the display (i.e. flickering etc).

5.
If there is any movement of signal, the extension wire must be taken apart and all connections checked and re-soldered. Then tested again until results are satisfactory.

6.
Check the cable is of correct quality standard. (See VM/COP/30.11 for details).

7.
Attach Hypertronics male 9 pin side to an Ohmeda monitor and the probe on to the ear to check SpO2 level. (Ideal reading 95-100.)

Labelling

1.
Labels: to be attached facing upwards as looking at the top of the probe.


· 1 x CE Label

· 1 x Green Ohmeda Label 


[image: image13.wmf]Green dumbell 

labell

CE

0086

Ø6 x 43 clear heat shrink

Serial No.

For use on

Ohmeda

Compatible

Pulse Oximeters


Quality Assurance (QA)

1.
Attach Hypertronics male 9-pin side to the test box connector marked ‘B’.

2.
Check display is showing correct characteristics as shown below. (At correct switch positions)


      LED should read approx 1.8v

[image: image14.jpg]


                    
[image: image15.jpg]R -

—— TN m—— T~

e

A AR e





            
     Position 1. IR, LED.
       
                           Position 4. Detector

3.
If the LED signal is at the bottom then it is wired incorrectly.

4.
‘Play’ with wire at connections to see if any change in the display (i.e. flickering etc).

5.
If there is any movement of signal, the extension wire must be taken apart and all connections checked and re-soldered. Then tested again until results are satisfactory.

6.
Check the cable is of correct quality standard. (See VM/COP/30.11 for details).

7.
Attach Hypertronics male 9-pin side to an Ohmeda monitor and the probe on to the ear to check SpO2 level. (Ideal reading 95-100.)

8.
Fill and sign attached paperwork.

Packaging 

1.
Visually check all labels are attached properly

2.
Using a twist tie (bunny clip) wrap the cable and place in a small blue Viamed plastic box, ensuring the cable is inserted in a neat and tidy presentable manor.

3.
Place a serial number sticker (supplied with the batch) on the front face of the box.

4.
Place a packed and tested sticker (also containing initials of the individual who is packing) on the right hand side top left corner of the box. Do not close box.

Final QA

1.
Final inspection. Visually ensure cable sit neatly within the box and is in a presentable state.

2.
Boxes are ready to stock in stores.
















Page 3 of 3

_1069495046.dwg

_1069504476.dwg

_1082896171.dwg

_1098013962.dwg

_1069495127.dwg

_1069494359.dwg

_1069494489.dwg

_1069494795.dwg

_1069494204.dwg

