[image: image11.png]


 

                           SpO2 ‘Y’ Probe Assembly

	COMPANY OPERATING PROCEDURES


0018692    
   Criticare   
         P869YA

VM3/COP/35.15
	Date:
14-May-02
	Revision date: 19-Sep-03
	Issue:   3




[image: image1.wmf]0.9m

D-Type Male 9 Pin

TOP

Y- Probe


[image: image2.jpg]






                     [image: image3.jpg]


             



Equipment required:
Soldering iron (0060120), solder (0050012), Wire stripper (0060030), Flush Cutter (0060010), Snipe nose pliers (0060021), ‘helping hand’ (0060145), Heat gun (0060100).

Parts list:
Kit and parts required. (Continued over page)

	D-type male 9-pin

Pre-manufactured
	Y Probe Side

Pre-manufactured

	Qty
	
	Part No.
	Qty
	Description
	Part No.

	
	
	
	
	
	

	1
	Pre- manufactured probes.
	0018692

	
	
	
	
	
	

	
	
	
	
	
	


ASSEMBLY OPERATIONS

1.
P869YA is a pre-manufactured unit that is not assembled by Viamed.

TESTING 

1.
Attach D-Type male 9-pin side to a CSI extension cable then to the test box connector marked  ‘J’.

2.
Check display is showing correct characteristics as shown below. (At correct switch positions)

       [image: image4.jpg]



               [image: image5.jpg]


                 [image: image6.jpg]




Pos 2. LED



  Pos 3. IR

                  Pos 4. Detector
3.
‘Play’ with wire at connections to see if any change in the display (i.e. flickering etc).

4.
If there is any movement of signal it must be reported to the supervisor. Then tested again until results are satisfactory.

5.
Check the cable is of correct quality standard. (See VM/COP/30.11 for details).

6.
Attach D-Type male 9-pin side to an extension cable then to a CSI 504 monitor and the probe on to the ear to check SpO2 level. (Ideal reading 95-100.)

Labelling

1.
Labels: to be attached facing upwards as looking at the top of the probe.


· 1 x CE Label

· 1 x Purple Criticare Label 


[image: image7.wmf]Purple dumbell 

labell

CE

0086

For use on

Criticare

Compatible

Pulse Oximeters


Quality Assurance (QA)

1.
Attach D-Type male 9-pin side to a CSI extension cable then to the test box connector marked  ‘J’.

2.
Check display is showing correct characteristics as shown below. (At correct switch positions)

       [image: image8.jpg]



               [image: image9.jpg]


                 [image: image10.jpg]




Pos 2. LED


       Pos 3. IR

         
         Pos 4. Detector
3.
‘Play’ with wire at connections to see if any change in the display (i.e. flickering etc).

4.
If there is any movement of signal, it must be reported to the supervisor. Then tested again until results are satisfactory.

5.
Check the cable is of correct quality standard. (See VM/COP/30.11 for details).

6.
Attach D-Type male 9-pin side to an extension cable then to a CSI 504 monitor and the probe on to the ear to check SpO2 level. (Ideal reading 95-100.)

7.
Fill and sign attached paperwork.

Packaging 

1.
Visually check all labels are attached properly

2.
Using a twist tie (bunny clip) wrap the cable and place in a small blue Viamed plastic box, ensuring the cable is inserted in a neat and tidy presentable manor.

3.
Place a serial number sticker (supplied with the batch) on the front face of the box.

4.
Place a packed and tested sticker (also containing initials of the individual who is packing) on the right hand side top left corner of the box. Do not close box.

Final QA

1.
Final inspection. Visually ensure cable sit neatly within the box and is in a presentable state.

2.
Boxes are ready to stock in stores.
















Page 2 of 3

[image: image11.png]_1082897379.dwg

_1083503570.dwg

