[image: image21.png]


 

                           SpO2 ‘Y’ Probe Assembly

	COMPANY OPERATING PROCEDURES


0018862    
    Kontron    
         P886YA

VM3/COP/35.26
	Date:
  22-May-02
	Revision date: 19-Sep-03
	Issue:   3




[image: image1.wmf]3.65m

Enlarged front view

Key-way Allignment

DPK1 Male 12 pin

TOP

Y- Probe


[image: image2.jpg]






                                   [image: image3.jpg]








Equipment required:
Soldering iron (0060120), solder (0050012), Wire stripper (0060030), Flush Cutter (0060010), Snipe nose pliers (0060021), ‘helping hand’ (0060145), Heat gun (0060100).

Parts list:
Kit and parts required. (Continued over page)

	SPK1-A male 12-pin Side
	‘Y’ Probe Side

	Qty
	Description
	Part No.
	Qty
	Description
	Part No.

	
	
	
	
	
	

	1
	SPK1-A male 10-pin kit
	0010707
	1
	Pre manufactured cable
	0018582

	(1)
	
[image: image4.wmf]       Rubber housing
	kit
	
	
	

	(1)
	      
[image: image5.wmf]                   Collar
	kit
	
	
	

	(5)
	
[image: image6.wmf]            Pins
	kit
	
	
	

	(1)
	 
[image: image7.wmf]        Upper casing
	kit
	
	
	

	(1)
	  
[image: image8.wmf]          Cable clamp
	kit
	
	
	

	(1)
	    
[image: image9.wmf]            Pin housing
	kit
	
	
	

	(2)
	      
[image: image10.wmf]                  Screws
	kit
	
	
	

	(1)
	
[image: image11.wmf]            Lower casing
	kit
	
	
	

	(1)
	      
[image: image12.wmf]             Maroon Collar
	kit
	
	
	

	1
	(6 x 43mm Clear heat shrink
	0032331
	
	
	

	1
	(6 x 25mm heat shrink
	0032321
	
	
	

	1
	(1.6 x 20mm heat shrink
	0032310
	
	
	

	1
	
[image: image13.wmf]R1

               3.3 k( Resistor
	0032070
	
	
	

	1
	
[image: image14.wmf]R1

               2.2 k( Resistor
	0032060
	
	
	


ASSEMBLY OPERATIONS

1.
Pre Heat soldering iron temperature to 240oc.

2.
Collect all required parts and equipment listed above.

‘Y’ Probe side:

1.
Probe side is pre-manufactured and ready to have the connector assembled to it.

SPK1-A male 12-pin Side :
Fig 2.2  
[image: image15.wmf]1

(-) Blue

Bottom Casing

Colour Collar

Pin Housing

Rubber Housing

Pins

Collar

Cable grip

Top Casing

(3) 

Combined 

outer and 

inner 

shields 

(is+ os) 

(7) White

(10) Yellow

(-) Orange

(3) Red

(12) Black

12

6

5

4

11

3

9

7

10

8

2

Rear View of Pin Housing

R2

R1

(3, 5) R1 = 3.3

W

K

(7, 8) R2 = 2.2

W

K


1.
Cut off the d-type connector from the end of the cable so to leave a cable length of 3m.

2.
Feed (6 x 43mm (clear) heat shrink, rubber housing, collar, cable grip and (6 x 25mm (black) heat shrink over the end of the cable.

3.
Strip 25mm off outer jacket of cable to reveal coloured wires, outer shield, and nylon/paper wire packing.

4.
Cut all packing orange and blue wire to the base. Trim resistor legs to 4mm and 15mm each.

5.
Strip 25mm off inner jacket to reveal black and white wires and the inner shield.

6.
Twist inner and outer shields together, and heat (1.6x 20mm heat shrink over excess naked wire.

7.
Trim ends of wires to the same length.

8.
Strip jacket of every wire 2mm to reveal copper core, and solder resistor legs and wires to the rear of individual pins.

9.
Insert all pins into correct locations.

10.
Insert pin housing into the bottom casing, and screw cable clamp over cable.

11.
Push top casing onto bottom casing, and push the collar over mating cases.

12.
Push rubber housing over cases, and then the maroon collar over the front of the casing.

TESTING 

1.
Attach SPK1-A side to a Kontron/Vickers box and then to the test box connector marked ‘A’.

2.
Check display is showing correct characteristics as shown below. (At correct switch positions)

[image: image21.png]
      LED should read approx 1.8v

[image: image16.jpg]


                    
[image: image17.jpg]R -

—— TN m—— T~

e

A AR e





            
     Position 1. IR, LED.
       
                           Position 4. Detector

3.
If the LED signal is at the bottom then it is wired incorrectly.

4.
‘Play’ with wire at connections to see if any change in the display (i.e. flickering etc).

5.
If there is any movement of signal, the cable must be taken apart and all connections checked and re-soldered. Then tested again until results are satisfactory.

6.
Check the cable is of correct quality standard. (See VM/COP/30.11 for details).

7.
Attach SPK1-A side to a Kontron monitor and the probe on to the ear to check SpO2 level. (Ideal reading 95-100.)

Labelling

1.
Labels: to be attached facing upwards as looking at the top of the probe.


· 1 x CE Label

· 1 x Serial no. Label (if required)

· 1x Orange ‘Do Not Throw Away’ Label.


[image: image18.wmf]0086

CE

Orange dumbell 

labell

Do Not

Throw

Away


Quality Assurance (QA)

1.
Attach SPK1-A side to a Kontron/Vickers box and then to the test box connector marked ‘A’.

2.
Check display is showing correct characteristics as shown below. (At correct switch positions)


      LED should read approx 1.8v

[image: image19.jpg]


                    
[image: image20.jpg]R -

—— TN m—— T~

e

A AR e





            
     Position 1. IR, LED.
       
                           Position 4. Detector

3.
If the LED signal is at the bottom then it is wired incorrectly.

4.
‘Play’ with wire at connections to see if any change in the display (i.e. flickering etc).

5.
If there is any movement of signal, the cable must be taken apart and all connections checked and re-soldered. Then tested again until results are satisfactory.

6.
Check the cable is of correct quality standard. (See VM/COP/30.11 for details).

7.
Attach SPK1-A side to a Kontron monitor and the probe on to the ear to check SpO2 level. (Ideal reading 95-100.)

8.
Fill and sign attached paperwork.

Packaging 

1.
Visually check all labels are attached properly

2.
Using a twist tie (bunny clip) wrap the cable and place in a small blue Viamed plastic box, ensuring the cable is inserted in a neat and tidy presentable manor.

3.
Place a serial number sticker (supplied with the batch) on the front face of the box.

4.
Place a packed and tested sticker (also containing initials of the individual who is packing) on the right hand side top left corner of the box. Do not close box.

Final QA

1.
Final inspection. Visually ensure cable sit neatly within the box and is in a presentable state.

2.
Boxes are ready to stock in stores.
















Page 4 of 4

_1070369243.dwg

_1077708075.dwg

_1083567074.dwg

_1083567913.dwg

_1083568454.dwg

_1077708209.dwg

_1070369391.dwg

_1070369566.dwg

_1070369328.dwg

_1070368994.dwg

_1070369122.dwg

_1070368914.dwg

_1069494204.dwg

