	Instructions for use: 

Series 400AC Temperature Probes

Contains: 1 piece 
	
	
	

	Applies to models ref:


2009/2010/2011/2012/2013/2014/2015/2016/2017/2018/2019/2020/2021/2022/2023/2024/2025/2026/2027/2028/2029/2030/
2031/2032   
	
	






























	Proper use

The Series 400AC reusable  temperature probes are intended for the continuous temperature measurement and monitoring of patients  when used with YSI, Hewlett Packard/Philips, Siemens and other manufacturers patient monitors and devices that are specially intended for use with YSI 400 series temperature probes and compatibles They are available in various configurations of sizes and types.
Cleaning and disinfecting

1. Before cleaning or disinfecting the probe, disconnect it from the monitoring device.

2. Clean and disinfect or autoclave the probe before using it on another patient.

3. Temperature probes can be cleaned with a cleaning agent by wiping the probe with a soft cloth . Rinse thoroughly with water.
4. To disinfect the probe, wipe the probe and the contact surfaces with disinfectant. For this we recommend the use of Cidex® OPA. If you intend to use a glutaric 

    aldehyde or an alternative product, please follow the usage instructions for these products. Rinse thoroughly with water.
Steam sterilization
Steam autoclave to 134°C using standard cycles. Follow the generally applicable instructions and regulations! Clean any residues from the temperature probe before sterilization.

Warning

1. This temperature probe is intended for use only in conjunction with compatible monitoring devices. Usage in conjunction with other devices can lead to malfunction.

2. Strong sources of electromagnetic fields, e.g. electro-surgical devices, can adversely affect the functioning of the equipment. If at all possible, the probe should be

    completely removed from the patient before a surgical device or  other HF source is activated! Check the protective earthing if you intend to use HF devices and 

    temperature measuring equipment in parallel. Minimise the potential danger by selecting a temperature measurement site distal  to the HF current path.

3. We do not recommend the use of the product within the range of Magnetic Resonance Imaging (MRI) or Computed Tomography (CT) devices.

4. Stop using the probe immediately if it becomes damaged.

5. For further information and warnings please read the usage instructions accompanying the temperature measuring device.

Caution

Improper use of temperature probes can lead to damage of the internal wiring and to loss of electrical insulation, as well as incorrect temperature readings.

Specifications 

Reusable and autoclavable  to 134 °C 

Storage temperature: -20°C…+70°C

Operating temperature:  +0°C...+50°C

Accuracy: 25°C…50°C ±0,1°C

Returning the product

Contact the customer service department at Viamed Ltd.  to obtain a RMA number (Return Material Authorization). Every product sent to us for checking must have been cleaned and disinfected beforehand.

[image: image2.wmf]0086

[image: image3.wmf]0086

[image: image4.jpg]


[image: image5.jpg]


[image: image1.jpg]VIAMED




VIAMED Ltd.

15 Station Road, Cross Hills, Keighley, 

West Yorkshire BD20 7DT, United Kingdom

Tel:    +44 (0)1535 634542

Fax:   +44 (0)1535 635582

Email: info@viamed.co.uk
www.viamed-online.com                                                                                                                                                                  Rev.4  valid from 10/07
[image: image6.jpg]



	
	


� EMBED CorelDraw.Graphic.8  ���








_1250407480.unknown

