[image: image14.png]


 

                          Temperature Probe Assembly

	COMPANY OPERATING PROCEDURES


0210401 Adult Rectal / Oesophageal

VM3/COP/34.20
	Date:
  7-Oct-03
	Revision date: 13-Oct-03
	Issue:   3




[image: image1.wmf]
  [image: image2.jpg]


                                                                          [image: image3.jpg]



Equipment required:
Soldering iron (0060120), solder (0050012), Wire stripper (0060030), Flush Cutter (0060010), Snipe nose pliers (0060021), ‘helping hand’ (0060145), Heat gun (0060100).

Parts list:
Kit and parts required.

	Parts

	Qty
	Description
	Part No.

	
	
	

	1
	
[image: image4.wmf]
Thermistor sub-assembly
	0231110

	1
	
[image: image5.wmf] Connector
	0231001

	1
	
[image: image6.wmf]                Casing
	0231001

	1
	
[image: image7.wmf]              Strain Relief
	0030618

	1
	
[image: image8.wmf]                       Cable Grip
	0030500

	1
	(6 x 30 mm Clear heat shrink
	0032331

	2
	(9 x 15mm Clear heat shrink
	0032332

	
	
	


ASSEMBLY OPERATIONS


[image: image9.wmf]Black

Shield

Cable grip

Strain relief

Casing

Connector


[image: image10.jpg]



1.
Pre Heat soldering iron temperature to 240oc.

2.
Collect all required parts and equipment listed above from the stores.

3.
Feed (6 x 30 mm clear heat shrink, casing, strain relief, and cable grip over cable end.

4.
Trim black wire to 15mm and shield wire to 8mm.

5.
Solder black wire to the top leg of the connector (see diagram) and heat the (9 x 15mm clear heat shrink over it to cover naked wire and bolt.

6.
Solder the shields (bare wire) to the correct surface and clamp the grip and cable grip around the cable. Heat the (9 x 15mm clear heat shrink over it to cover naked metal.

7.
Screw the casing onto the connector.

Labelling

1.
Labels to be attached:


· 1 x CE Label

· 1 x Serial no. Label

· [image: image14.png]1x White ‘do not throw away’ dumbbell label with the following label attached to the reverse side:


[image: image11.wmf]CE

0086

Ø6 x 30 clear heat shrink

Serial No.

Adult rectal/esophageal

White dumbell 

labell

37104

temperatue probe


Testing

1.
Using the test equipment check to see if the probe works. Reading for mono should show between 2 and 3 k(.

[image: image12.jpg]


      [image: image13.jpg]



Packaging and Storage

1.
Visually check all labels are attached properly.

2.
Using a twist tie (bunny clip) wrap the cable and place in a plastic bag. Attach a label on the bag showing all necessary details. Also seal the bag using a packed and tested sticker.


3.
Place 5 bagged units into a larger plastic bag (pt no. 9950211) and book into stores.

37104


Adult rectal/oesophageal


temperature probe











37104


Adult rectal/oesophageal


temperature probe








Page 3 of 3

_1090848352.dwg

_1094295151.dwg

_1090913969.dwg

_1090837400.dwg

_1090837439.dwg

_1090841408.dwg

_1090837323.dwg

_1090837367.dwg

