[image: image15.png]


 

                Temperature Probe Assembly

	COMPANY OPERATING PROCEDURES


0210403 Paediatric Surface / Skin

VM3/COP/34.30
	Date:
  7-Oct-03
	Revision date: 13-Oct-03
	Issue:   3




[image: image16.jpg]



[image: image2.jpg]


              

                 [image: image3.jpg]



Equipment required:
Soldering iron (0060120), solder (0050012), Wire stripper (0060030), Flush Cutter (0060010), Snipe nose pliers (0060021), ‘helping hand’ (0060145), Heat gun (0060100).

Parts list:
Kit and parts required.

	Parts

	Qty
	Description
	Part No.

	
	
	

	1
	
[image: image4.wmf] Thermistor
	0231100

	1
	
[image: image5.wmf]     Connector
	0231001

	1
	
[image: image6.wmf]                     Casing
	0231001

	1
	
[image: image7.wmf]              Strain Relief
	0030618

	1
	
[image: image8.wmf]                       Cable Grip
	0030500

	1
	(6 x 40 mm Clear heat shrink
	0032331

	2
	(9 x 15mm Clear heat shrink
	0032332

	1
	(6 x 40 mm Black heat shrink
	0032321

	1
	Can of clear acrylic spray
	

	
	
	


ASSEMBLY OPERATIONS


[image: image9.wmf]Grey

Grey

Cable grip

Strain relief

Casing

Connector


[image: image10.jpg]



1.
Pre Heat soldering iron temperature to 240oc.

2.
Collect all required parts and equipment listed above from the stores.

3.
Feed (6 x 30 mm clear heat shrink, casing, strain relief, and cable grip over cable end.

4.
Heat (6 x 40 black heat shrink over the grey wires leaving approximately 8mm of grey wire visible at the end and strip 2mm off the ends.

5.
Solder one grey wire to the leg of the connector (see diagram) and heat the (9 x 15mm clear heat shrink over it to cover naked wire and bolt.

6.
Solder the other grey wire to the surface and clamp the grip and cable grip around the cable. Heat the (9 x 15mm clear heat shrink over it to cover naked metal.

7.
Screw the casing onto the connector.
Labelling

1.
Labels to be attached:


· 1 x CE Label

· 1 x Serial no. Label

· 1x White ‘do not throw away’ dumbbell label with the following label attached to the reverse side:

[image: image1.wmf]

[image: image11.wmf]CF60901895

Serial No.

0086

CE

Ø6 x 30 clear heat shrink

White dumbell 

labell

Probe. Surface/Skin - 37804

0210403

Neonatal Temperature


Re-Spraying

1.
Cover area of wire leaving just the thermistor head showing (10 at a time).

[image: image12.jpg]



2.
Take all covered units to the spraying booth and spray the heads.

3.
Leave to dry for 1 hour and repeat to give 4 to 5 coats (or until a good quality of finish can be seen).








Acceptable finish

[image: image15.png]
          Unacceptable finish

Testing

1.
Using the test equipment check to see if the probe works. Reading for mono should show between 2 and 3 k(.

[image: image13.jpg]


  [image: image14.jpg]



Packaging and Storage

1.
Visually check all labels are attached properly.

2.
Using a twist tie (bunny clip) wrap the cable and place in a plastic bag. Attach a label on the bag showing all necessary details. Also seal the bag using a packed and tested sticker.


3.
Place 10 bagged units into a larger plastic bag (pt no. 9950211) and book into stores.

0210403


Paediatric Temperature probe. Surface/Skin –37804


CF60901895





0210403


Paediatric Temperature probe. Surface/Skin –37804


CF60901895









Page 1 of 4

_1094280071.dwg

_1094280127.dwg

_1090837400.dwg

_1090924123.dwg

_1090924389.dwg

_1090837439.dwg

_1090837323.dwg

_1090837367.dwg

