[image: image15.png]


 

                          Temperature Probe Assembly

	COMPANY OPERATING PROCEDURES


0210701 Adult Rectal

VM3/COP/34.50
	Date:
7-Oct-03
	Revision date: 8-Dec-03
	Issue:   3




[image: image1.wmf]
  [image: image2.jpg]


                                                       [image: image3.jpg]



Equipment required:
Soldering iron (0060120), solder (0050012), Wire stripper (0060030), Flush Cutter (0060010), Snipe nose pliers (0060021), ‘helping hand’ (0060145), Heat gun (0060100).

Parts list:
Kit and parts required.

	Parts

	Qty
	Description
	Part No.

	
	
	

	1
	
[image: image4.wmf]
Thermistor sub-assembly
	0231125

	1
	
[image: image5.wmf]       Connector
	0231000

	1
	
[image: image6.wmf]                Casing
	0231000

	1
	
[image: image7.wmf]              Strain Relief
	0030618

	1
	(6 x 30 mm Clear heat shrink
	0032331

	2
	(9 x 15mm Clear heat shrink
	0032332

	
	
	


ASSEMBLY OPERATIONS

Schematic view only


[image: image8.wmf]Black

Red

Cable grip

Strain relief

Casing

Connector

White


1.
Pre Heat soldering iron temperature to 240oc.

2.
Collect all required parts and equipment listed above from the stores.

3.
Feed (6 x 30 mm clear heat shrink, casing, strain relief, and cable grip over cable end.

4.
Trim wires to 20mm and place blue cover over red and black wires, and the clear heat-shrink over the black wire.

[image: image9.jpg]



[image: image10.jpg]



5.
Solder the white wire to the long base, black wire to the top left (small leg), and red wire to the top right (long leg) of the connector (see picture above) and heat the (9 x 15mm clear heat shrink over the black wire to cover the connection and bolt. Then push the blue cover over the red and black wires.

[image: image11.jpg]



6.
Clamp the outer casing with the cable grip.

7.
Screw the cover on to the connector.

Labelling

1.
Labels to be attached:


· 1 x CE Label

· 1 x Serial no. Label

· 1x White ‘do not throw away’ dumbbell label with the following label attached to the reverse side:

[image: image15.png]

[image: image12.wmf]Serial No.

0086

CE

Ø6 x 30 clear heat shrink

White dumbell 

labell

37107

0210701

Adult rectal/Oesophageal

CI61673601


Testing

1.
Using the test equipment check to see if the probe works. Reading for stereo should show 6 and 30 k(. (2 switch positions for 2 readings).

[image: image13.jpg]




[image: image14.jpg]



Packaging and Storage

1.
Visually check all labels are attached properly.

2.
Using a twist tie (bunny clip) wrap the cable and place in a plastic bag. Attach a label on the bag showing all necessary details. Also seal the bag using a packed and tested sticker.


0210701


Adult Rectal/Oesophageal


37107


CI61673601











0210701


Adult Rectal/Oesophageal


37107


CI61673601








Page 3 of 5

_1094304017.dwg

_1094364178.dwg

_1127292605.dwg

_1094304123.dwg

_1090837439.dwg

_1094295151.dwg

_1090837400.dwg

