Apgar  Timer Risk Report


Pre-design considerations

Power suppl;y

The unit is to be used in the resuscitation cabinet.

100% Oxygen is fed  into the Cabinet from wall outlets 

This creates a risk if Oxygen leaks whilst the unit is on and is increased if the cabinet door is closed. Battery versions were therefore chosen.

Low voltage less than 6  volt was preferred.

Function

Display:


Options of 




Real time clock




Elapsed Time clock




Combination




Alarm Indication only

Display Size






Should be large enough to be read at 1 metre




Should inform battery status




Backlight not required

Real Time clock: 



possible risk of time not in sync with other visible clocks in the room 

causing confusion

LED Alarm only




Adequate 

Alarms




Audio :loudness controlled by ISO standards




Frequency : 1 minute: 5 minutes: 10 minutes

Controls:

Start




Pause ( for reading only) timer is not stopped




Reset: needs to be push & Hold or multiple push to prevent 



accidental stop

Page 

1 December 2004


