[image: image1.png]


[image: image2.jpg]ARAB HEALTH
2002

Summary of products

Pulse Oximetry Sensors Viamed  Dolphin  Aristo
Finger sensors v v v
Multi-site 'Y’ sensors v v v
Ear sensors v
Extension cables v v v
Adapter cables v v v
D.O.T. Sensors v
Wrap sensors v
Disposables v v ~

Pulse oximetry sensor repair services
Pulse oximeters - Stand alone, hand held, compact PC
Pulse oximetry simulators and testers

Oxygen sensors
Oxygen monitors and analysers

Oxygen hoods
Patient monitors
Infant resuscitation cabinets

Infant resuscitators
Radiant warmers

Phototherapy masks
Phototherapy light shields

ECG cables

Temperature probes

Supra-maximal nerve stimulator

Chart & recorder paper [
VIARMED

Imail infeid





                       

                              APGAR TIMER - CE FILE.

Risk Assessment iaw EN ISO 14971:2000 Annex D : Possible hazards with medical devices.

	Ref.
	Hazard.

	Related part / Component posing risk.
	Sev

of

Haz.
	Like

of

Haz.
	Risk.
	Solution.
	Document referenced.
	Sev

of

Haz.
	Like

0f

Haz.
	Risk.

	D.2
	Energy hazards and contributory factors

	D.2.1
	Electricity
	N/A
	1
	1
	1
	
	(YZ) Design Specification
	1
	1
	1

	D.2.2
	Heat
	N/A
	1
	1
	1
	
	
	1
	1
	1

	D.2.3
	Mechanical force
	Instrument
	2
	1
	2
	(“May be damaged if dropped” in User Instructions)
	(F) User manual 
	1
	1
	1

	D.2.4
	Ionising radiation
	N/A
	1
	1
	1
	No Radiation
	
	1
	1
	1

	D.2.5
	Non ionising radiation
	N/A
	1
	1
	1
	No Radiation
	
	1
	1
	1

	D.2.6
	Moving parts
	N/A
	1
	1
	1
	No Moving Parts
	
	1
	1
	1

	D.2.7
	Unintended motion
	N/A
	1
	1
	1
	Fixed
	
	1
	1
	1

	D.2.8
	Suspended masses
	N/A
	1
	1
	1
	None
	
	1
	1
	1

	D.2.9
	Patient support failure
	N/A
	1
	1
	1
	Not Applicable
	
	1
	1
	1

	D.2.10
	Pressure (vessel rupture)
	N/A
	1
	1
	1
	Not Applicable
	
	1
	1
	1

	D.2.11
	Acoustic pressure
	N/A
	1
	1
	1
	Not Applicable
	
	1
	1
	1

	D.2.12
	Vibration
	N/A
	1
	1
	1
	Not Applicable
	
	1
	1
	1

	D.2.13
	Magnetic fields 

(eg. MRI)
	N/A
	1
	1
	1
	Do not use in MRI
	(F) User manual 
	1
	1
	1

	D.3
	Biological hazards and contributory factors

	D.3.1
	Bio-contamination
	Instrument
	2
	2
	4
	Easy to clean – cleaning instructions in user manual
	(F) User manual 
	1
	1
	1

	D.3.2
	Bio-incompatibility
	Instrument
	1
	1
	1
	No adverse materials used
	Manufacturers data
	1
	1
	1

	D.3.3
	Incorrect formulation (chemical composition)
	N/A
	1
	1
	1
	Not Applicable
	
	1
	1
	1

	D.3.4
	Toxicity
	N/A
	1
	1
	1
	Not Applicable
	Manufacturers data
	1
	1
	1

	D.3.5
	Allergenicity
	N/A
	1
	1
	1
	Not Applicable
	Manufacturers data
	1
	1
	1

	D.3.6
	Mutagenicity
	N/A
	1
	1
	1
	Not Applicable
	Manufacturers data
	1
	1
	1

	D.3.7
	Oncogenicity
	N/A
	1
	1
	1
	Not Applicable
	Manufacturers data
	1
	1
	1

	D.3.8
	Carcinogenicity
	N/A
	1
	1
	1
	Not Applicable
	Manufacturers data
	1
	1
	1

	D.3.9
	Re and/or cross infection
	Instrument & leads
	1
	1
	1
	Not Applicable
	Manufacturers data
	1
	1
	1

	D.3.10
	Pyrogenicity
	N/A
	1
	1
	1
	Not Applicable
	Manufacturers data
	1
	1
	1

	D.3.11
	Inability to maintain hygienic standards
	Instrument 
	2
	2
	4
	Easy to clean – cleaning instructions in user manual
	(F) User manual 
	1
	1
	1

	D.3.12
	Degradation
	Instrument 
	2
	1
	2
	Care instructions given in the user manual
	(F) User manual 
	1
	1
	1

	D.4
	Environmental hazards and contributory factors

	D.4.1
	Electromagnetic fields
	N/A
	1
	1
	1
	Instrument Design
	
	1
	1
	1

	D.4.2
	Susceptibility to electromagnetic interference
	N/A
	1
	1
	1
	Instrument Design
	
	1
	1
	1

	D.4.3
	Emissions of electromagnetic interference 
	N/A
	1
	1
	1
	Instrument Design
	
	1
	1
	1

	D.4.4
	Inadequate supply of power 
	N/A
	1
	1
	1
	Battery Only
	(YZ) Design Files
	1
	1
	1

	D.4.5
	Inadequate supply of coolant
	N/A
	1
	1
	1
	No Coolant
	
	1
	1
	1

	D.4.6
	Storage / operation outside prescribed environmental conditions
	Instrument
	2
	1
	2
	No specific environmental storage / operating conditions imposed
	
	1
	1
	1

	D.4.7
	Incompatibility with other devices with which the product is intended to be used
	Instrument
	1
	1
	1
	Battery - Stand alone
	
	1
	1
	1

	D.4.8
	Accidental mechanical damage
	Instrument 
	2
	1
	2
	“May be damaged if dropped” in User Instructions. 
	(F) User manual 
	1
	1
	1

	D.4.9
	Contamination due to waste products and/or device disposal
	Instrument 
	1
	1
	1
	Electronic waste disposal required. RoHS
	
	1
	1
	1

	D.5
	Hazards resulting from incorrect output of energy and substances

	D.5.1
	Electricity
	N/A
	1
	1
	1
	Not Applicable
	
	1
	1
	1

	D.5.2
	Radiation
	N/A
	1
	1
	1
	Not Applicable
	
	1
	1
	1

	D.5.3
	Volume
	N/A
	1
	1
	1
	Not Applicable
	
	1
	1
	1

	D.5.4
	Pressure
	N/A
	1
	1
	1
	Not Applicable
	
	1
	1
	1

	D.5.5
	Supply of medical gases
	N/A
	1
	1
	1
	Not Applicable
	
	1
	1
	1

	D.5.6
	Supply of anaesthetic agents
	N/A
	1
	1
	1
	Not Applicable
	
	1
	1
	1

	D.6
	Hazards related to the use of the medical device and contributory factors

	D.6.1
	Inadequate labelling
	Label
	2
	1
	2
	No instruction label on instrument – Instructions for use attached
	(F) User manual 
	1
	1
	1

	D.6.2
	Inadequate operating instructions
	Instructions for Use
	2
	1
	2
	Instructions for use comprehensively cover the product.  Instrument only for use by trained and qualified personnel. 
	(F) User manual 
	1
	1
	1

	D.6.3
	Inadequate specification of accessories
	N/A
	1
	1
	1
	Not Applicable
	
	1
	1
	1

	D.6.4
	Inadequate specification of pre-use checks
	Instructions for Use
	2
	2
	4
	Instrument must be pre-use checked – Instructions for Use refer. 
	(F) User manual 
	2
	1
	2

	D.6.5
	Over-complicated operating instructions
	Instructions for Use
	2
	1
	2
	Instructions for use 
	(F) User manual 
	1
	1
	1

	D.6.6
	Inadequate specification of service and maintenance
	Instructions for Use


	1
	1
	1
	No service or maintenance except batteries
	
	1
	1
	1

	D.6.7
	Use by unskilled / untrained personnel
	Instrument 
	2
	2
	4
	
	(F) User manual 
	1
	1
	1

	D.6.8
	Reasonable foreseeable misuse
	Instrument 
	1
	1
	1
	Batteries removed for long periods of non-use
	
	2
	1
	2

	D.6.9
	Insufficient warning of side effects
	N/A
	1
	1
	1
	Not Applicable
	
	1
	1
	1

	D.6.10
	Inadequate warnings of hazards likely with re-use of single use devices
	N/A
	1
	1
	1
	Not Applicable
	
	1
	1
	1

	D.6.11
	Incorrect measurement and other metrological aspects
	N/A
	1
	1
	1
	Not Applicable
	
	1
	1
	1

	D.6.12
	Misrepresentation of results
	N/A
	1
	1
	1
	Not Applicable
	
	1
	1
	1

	D.6.13
	Incompatibility with consumables / accessories / other devices
	N/A
	1
	1
	1
	Not Applicable
	
	1
	1
	1

	D.6.14
	Sharp edges or points
	Instrument Case
	1
	1
	1
	
	(F) User manual 
	1
	1
	1

	D.7
	Inappropriate, inadequate or overcomplicated user interface ( man/machine communication)

	D.7.1
	Mistakes & judgement errors
	Instrument
	3
	1
	3
	Only for use by qualified & trained personnel – warnings in user / technical manual. Pre-use check always before use.
	(F) User manual 
	1
	1
	1

	D.7.2
	Lapses and cognitive recall errors
	Instrument
	3
	1
	3
	Only for use by qualified & trained personnel – warnings in user / technical manual. Preuse check always before use.
	(F) User manual 
	1
	1
	1

	D.7.3
	Slips & blunders (mental or physical)
	Instrument
	3
	1
	3
	Only for use by qualified & trained personnel – warnings in user / technical manual. Preuse check always before use.
	(F) User manual 
	1
	1
	1

	D.7.4
	Violation or abbreviation of instructions,  procedures etc
	Instrument
	3
	1
	3
	Only for use by qualified & trained personnel – warnings in user / technical manual. Preuse check always before use.
	(F) User manual 
	1
	1
	1

	D.7.5
	Complex or confusing control system
	N/A
	1
	1
	1
	No complex user interface
	
	1
	1
	1

	D.7.6
	Ambiguous or unclear device state
	Instrument
	3
	1
	3
	Only for use by qualified & trained personnel – warnings in user / technical manual. Preuse check always if device state unclear.
	(F) User manual 
	1
	1
	1

	D.7.7
	Ambiguous or unclear presentation of settings, measurement, or other information
	N/A
	1
	1
	1
	No complex user interface
	International Symbols
	1
	1
	1

	D.7.8
	Misrepresentation of results
	Instrument
	3
	1
	3
	Only for use by qualified & trained personnel – warnings in user / technical manual.
	(F) User manual 
	1
	1
	1

	D.7.9
	Insufficient visibility, audibility or tactility
	N/A
	1
	1
	1
	No complex user interface
	
	1
	1
	1

	D.7.10
	Poor mapping of controls to action or of displayed information to actual state
	N/A
	1
	1
	1
	Only for use by qualified & trained personnel – warnings in user / technical manual.
	(F) User manual 
	1
	1
	1

	D.7.11
	Controversial modes or mappings as compared to existing equipment
	N/A
	1
	1
	1
	No complex user interface
	
	1
	1
	1

	D.8
	Hazards arising from function failure, maintenance and ageing and contributory factors

	D.8.1
	Erroneous data transfer
	N/A
	1
	1
	1
	Not Applicable
	
	1
	1
	1

	D.8.2
	Lack of, or inadequate specification for maintenance including post maintenance functional tests
	Instrument
	1
	1
	1
	Instructions for use state service requirements. Instrument will not pass pre-use check if badly damaged.
	(F) User manual 
	1
	1
	1

	D.8.4
	Inadequate maintenance
	Instrument
	1
	1
	1
	Instrument will not pass pre-use check if servicing required.
	(F) User manual 
	1
	1
	1

	D.8.5
	Lack of adequate determination of end of device life
	Instrument
	1
	1
	1
	User decision based on damage or cost effectiveness of repair. Instrument will not pass preuse check if badly damaged or requiring servicing.
	
	1
	1
	1

	D.8.6
	Loss of electrical integrity
	N/A
	1
	1
	1
	Not Applicable
	
	1
	1
	1

	D.8.7
	Loss of mechanical integrity
	Instrument
	1
	1
	1
	User decision based on damage or cost effectiveness of repair.  Instrument will not pass pre-use check if badly damaged.
	
	1
	1
	1

	D.8.8
	Inadequate packaging (contamination and / or deterioration of the device)
	Instrument 
	1
	1
	1
	Bubble wrapped and boxed prior to despatch.
	(M) Packaging


	1
	1
	1

	D.8.9
	Re-use and/or improper re-use
	Instrument 
	1
	1
	1
	User decision on suitability for next use based on pre use check & tolerable mechanical damage.
	 
	1
	1
	1

	D.8.10
	Deterioration in function (gradual occlusion of fluid / gas path or change in resistance to flow, electrical conductivity) as a result of repeated use
	Instrument 
	1
	1
	1
	Not Applicable
	
	1
	1
	1



Severity of hazard: 1 = Insignificant, 2 = Tolerable, 3  = Critical, 4  = Intolerable.  Probability of event: 1 = Improbable, 2 = Occasional, 3 = Likely, 4 = Highly likely.  

Risk calculated as severity of hazard x probability of event, 1 – 16.  Further explanation of risk management policy – see Risk Management Policy & Definitions.

E. Risk Management BS14971 Annex D_Apgar Timer
16/12/2004
    




Page 7

