APGAR TIMER

The Viamed Apgar Timer is a battery powered digital timer with preset audible timing indications to assist with the Apgar Scoring of neonates.

Standby Mode

The Apgar Timer does not have an On/Off switch. 

When not counting, the Apgar Timer enters Standby Mode. 

In Standby Mode the display is blank with the exception of a flashing colon to indicate the timer is functioning correctly.

If the display is completely blank with no flashing colon, the timer may require battery replacement.

Counting Mode

When the Apgar Timer begins counting, the display counts upwards from 00:00 in 1-second intervals up to a maximum of 1 hour (59:59 on display), it will then enter Standby Mode.

Whilst counting, the timer will indicate the Apgar Scoring intervals with beeps:

· 1 beep:

1 minute 


· 2 beeps:

5 minutes

· 3 beeps:

10 minutes

Freeze Mode

It is possible to freeze the display whilst the actual time continues to count in the background; this is to allow for recording the time from the start of the counting cycle of a particular event. 

When the display is unfrozen the timer will display the actual elapsed time and continue counting.

Should an Apgar Scoring interval be reached whilst the display is frozen, then the timer will still beep to indicate that interval but will remain in Freeze Mode until exited by the user.

Reset Function

The Reset Function allows the timer to be reset and placed back into Standby Mode.
During the reset procedure, a self-test function tests every segment of the LCD (8.8:8.8 on the display), if there are missing segments the Apgar Timer should be returned to Viamed for repair.

Low Battery Indications

The LCD flashing between the messages “Lo” and “Batt” indicates a low battery condition.

Additionally, 3 decimal points on the display indicate the approximate level of battery life remaining, with 3 points reducing to 2, and finally 1 point as the battery life diminishes.

This ‘life remaining’ indicator is for indication only and may not correspond directly to a linear decay of the remaining battery life.

The Functions of the Keys

	Key


	Function

	[image: image1.jpg]




	Starts counting from 00:00

Returns to Counting Mode whilst in Freeze Mode

	[image: image2.jpg]




	Enters and exits Freeze Mode
Display ‘freezes’ whilst the actual elapsed time continues counting but is not displayed, exiting Freeze Mode returns to Counting Mode at the actual elapsed time

	[image: image3.jpg]</





	Resets counter, returning timer to Standby Mode
During reset, the display will briefly show 8.8:8.8 to test the LCD segments


Battery Replacement

-Continue from here

Type of batteries 1.5V AA / LR6 / MN1500 Alkaline

Qty 4

