[image: image1.png]


 

                   O2 Sensors

	COMPANY OPERATING PROCEDURES


Teledyne K1 (PN 0110210)
O2 Sensor Test Procedure

VM3/COP/37.10
	Date:
  3-Feb-03
	Revision date: 28-Jul-06
	Issue:   3



Test & Equipment Required

Visual check – Digital Voltmeter

Output specification as in document: (VM3COP37.00, column (i)).

Method.

1. Open the packaging and remove the O2 sensor.

2. Check the O2 sensor for damage and signs of leaking electrolyte.

3. Check output in Air  10mV +/- 3mV 

4. Label with Viamed stickers if appropriate.

5. Reseal in serial numbered packet, integrity seal the edge of the packet with initialed and dated Viamed sticker and re-box. 
















VM3COP 37.10 K1.doc                  28/07/2006                                Page 1 of 1

[image: image1.png]