[image: image1.jpg]ARAB HEALTH
2002

Summary of products

Pulse Oximetry Sensors Viamed  Dolphin  Aristo
Finger sensors v v v
Multi-site 'Y’ sensors v v v
Ear sensors v
Extension cables v v v
Adapter cables v v v
D.O.T. Sensors v
Wrap sensors v
Disposables v v ~

Pulse oximetry sensor repair services
Pulse oximeters - Stand alone, hand held, compact PC
Pulse oximetry simulators and testers

Oxygen sensors
Oxygen monitors and analysers

Oxygen hoods
Patient monitors
Infant resuscitation cabinets

Infant resuscitators
Radiant warmers

Phototherapy masks
Phototherapy light shields

ECG cables

Temperature probes

Supra-maximal nerve stimulator

Chart & recorder paper [
VIARMED

Imail infeid




[image: image2.png]



                                          APGAR TIMER - CE File

Modification - Design History

2002.

The original Laboratory Analogue Clock, being no longer available, is to be replaced with a Digital version for the “Stop / Start / Reset” functions.

A search of all, watch and clock manufacturers, has found that the nearest available for the required function has too many functions and could therefore cause confusion and potential delay in the operation of the Tom Thumb unit.

This was deemed to be unacceptable, therefore the clock with the functions closest to the requirements was purchased with a view to a re-design / modification, to reduce the functions to the basic three required:

Start

Stop

Reset

This clock was purchased from E.A. Combs of London.

Modifications would include the removal of buttons no longer required, plus the removal of the three jack plug sockets.

Replacement of the “Front-case” label, with a bespoke Viamed Label.

Incorporation of a “Single” restart switch.

Addition of a Dovetail attachment for “Rail Clamping” of the complete unit to the Resuscitation Cabinet.

T:\Kevin Rush\CE Files\CE files current\Apgar Timer\Apgar Timer_CE File_2004\Apgar Timer_Section O_Design Files_2004\O. 01_01_Design History.doc


26/01/2004                                                                                                                                             Page 1



[image: image1.jpg][image: image2.png]