[image: image1.png]


 

                   O2 Sensors

	COMPANY OPERATING PROCEDURES


Maxtec Max-10 (PN 0110410)
O2 Sensor Test Procedure

VM3/COP/37.16
	Date:
  3-Feb-03
	Revision date: 21-Apr-05
	Issue:   3



Test & Equipment Required

Functional test required using MAX-10 tester & adapter cable (digital multimeter & adapter cable).

Output specification as in document: (VM3COP37.00, column (i)).

Method.

1. Open the packaging and remove the O2 sensor.

2. Check the O2 sensor for damage and signs of leaking electrolyte.

3. Check for an output

4. If no, or very low output, or sensor is returned as faulty leave the O2 sensor exposed to air for 2 Hours

5. Connect the O2 sensor to the MAX-10 tester, select 200mV range (marked with arrows) and wait for meter reading to stabilise.

6. Ensure that meter reading is within the output specification range, as in VM3COP37.00, column (i). If the O2 sensor fails the output specification test, it should be resealed and boxed, placed in a docket and put on the goods in desk with a short note explaining why it failed, ready to be booked in.

7. Disconnect the O2 sensor from the MAX-10 tester and label with Viamed stickers if appropriate.

8. Reseal in serial numbered packet, integrity seal the edge of the packet with initialed and dated Viamed sticker and re-box.

9. Ensure the O2 sensor is booked out in the stock books.















Page 1 of 1

[image: image1.png]