APGAR Timer Power Consumption.

Author:
Peter Anderson

Date:
21st October 2004

Assumptions:

1
The timer will never be turned OFF – there will always be some LCD activity.

2
The timer will be used twice a day, 365 days a year.

3
The sounder will operate at the 1, 5 and 10 minute points.  Each operation will be for 5 seconds at a 50/50 duty cycle.

4
The LCD will be updated every second, 365 days a year.

5
The battery will be monitored once every hour with 5 ADC cycles per sample.

Power Consumption Calculation:


Number of hours in a year:
24 x 365





8760

ST7LITE15


CPU (Estimated):



Active Halt mode
200uA



Run (Slow) mode
200uA



Consumption:
200uA x 8760




1752 mAh

(1)


LVD/AVD


245uA





245uA x 8760




2146 mAh

(2)


RC Oscillator (Not used)






0 mAh


PLL (Not used)







0 mAh


12-bit Timer AT2

50uA





50uA x 2(Use) x (3(Times) x 5 (Seconds) / 3600(Hours) x 365) / 2(Duty)
0.076 mAh

(3)


SPI (Not used)







0 mAh


ADC

1.2mA





1.2mA x 5(Samples) x 0.5mS(Conversion time) / 3600(Hours) x 8760
0.0073 mAh
(4)


Flash Memory



No Read/Write
100uA





100uA x 8760




876 mAh

(5)



Read

2.6mA





2.6mA x (256 / 32768)(Instructions) x 8760


178 mAh

(6)


Ports (Leakage)

1uA





1uA x 15(Pins) x 8760




131 mAh

(7)

PCF8576D LCD Driver


Supply

20uA





20uA x 8760




175 mAh

(8)


LCD Supply

60uA





60uA x 8760




526 mAh

(9)

Sounder


2.5mA





2.5mA x 2(Use) x (3(Times) x 5 (Seconds) / 3600(Hours) x 365) / 2(Duty)
3.8 mAh

(10)

I2C interface

2(Lines) x 3V / 2700 Ohms x 1 / 1000(duration) / 3600(Hours) x 8760
0.00 mAh

(11)

Battery Potential Divider

10uA





10uA x 8760




88 mAh

(12)

LTC1877 Regulator (1/5mA Load)
100uA




100uA x 8760




876 mAh

(13)








Total (1 – 13):
6752 mAh

Battery capacity required:
6752 mAh

“AA” Battery capacities by manufacturer:

Varta 4106


2600 mAh

Duracell MN1500

2700 mAh

Energizer Ultra Alkaline

2850 mAh

Number of batteries required:
6752 / 2600





3 batteries

