Instructions for use: 

Series 400 & Series 700 Temperature Probes

Contains: 1 piece
Applies to models ref:

37104, 37204, 37804, 37904, 37107, 37907 37104/D, 37204/D, 37804/D, 37904/D
See overleaf for details.

Proper use

The Series 400 & Series 700 reusable temperature probes are intended for the continuous temperature measurement and monitoring of patients when used with YSI patient monitors and devices, that are specially intended for use with YSI 400 & YSI 700 series temperature probes and compatibles and Dolphin Medical pulse oximeters. They are available in various configurations of sizes and types.

Instructions for use

Temperatures can be taken by placing the temperature probe on the skin surface, in the oesophagus, or rectally using the appropriate model.

Skin

1. Position the probe with a metal side to the skin.

2. Secure the probe using a heat reflective pad, taking care not to compress blood vessels.

3. Secure the lead using medical adhesive tape. 
4. The temperature probe requires a three to five minute stabilising time.
Oesophageal
1. Introduce the probe until it has reached the lower third of the oesophagus. 
2. Secure the lead using medical adhesive tape.
3. The temperature probe requires a three to five minute stabilising time.
Rectal

1. Place a disposable rectal probe cover over the end of the temperature.
2. Lubricate the end of the probe.
3. Insert the probe into the rectum.
4. The probe is in the correct position when it has passed the sphincter.

5. Secure the lead using medical adhesive tape.
6. The temperature probe requires a three to five minute stabilising time.
Cleaning and disinfecting

1. Clean and disinfect before using the temperature probe on another patient.

2. Before cleaning or disinfecting the probe, disconnect it from the monitoring device.

3. Clean the probe of any adhesive tape and bioburden material prior to disinfection or sterilisation.

4. Temperature probes can be cleaned with a cleaning agent by wiping the probe with a soft cloth.

5. To wipe clean, hold the lead near the probe end and wipe towards the plug.
6. To disinfect, wipe the probe and contact surfaces with disinfectant. For this we recommend the use of CIdex® OPA. 

7. If you intend to use a glutaric aldehyde or an alternative product, please follow the usage instructions for these products. 

8. Rinse thoroughly with water.
Specifications 

Reusable 

Accuracy: +/- 0.1 ˚C
Warning

1. This temperature probe is intended for use only in conjunction with compatible monitoring devices. Usage in conjunction with other devices can lead to malfunction.

2. Strong sources of electromagnetic fields, e.g. electro-surgical devices, can adversely affect the functioning of the equipment. If at all possible, the probe should be completely removed from the patient before a surgical device or other HF source is activated! Check the protective earthing if you intend to use HF devices and temperature measuring equipment in parallel. Minimise the potential danger by selecting a temperature measurement site distal to the HF current path.

3. We do not recommend the use of the product within the range of Magnetic Resonance Imaging (MRI) devices.

4. Stop using the probe immediately if it becomes damaged.

5. For further information and warnings please read the usage instructions accompanying the temperature measuring device.

Caution

Improper use of temperature probes can lead to damage of the internal wiring and to loss of electrical insulation, as well as incorrect temperature readings.

Returning the product

Contact the customer service department at Viamed Ltd.  to obtain a RMA number (Return Material Authorization). Every product sent to us for checking must be cleaned and disinfected beforehand.
[image: image1.jpg]VIAMED



[image: image2.wmf]0086


[image: image3.wmf]0086

[image: image4.jpg]



	Part Number
	Model Ref.
	Description
	Compatibility 
	Length

(metres)
	Connector

	0210401
	37104
	Adult rectal/oesophageal.
	YSI 400 series.
	3
	6.35mm (¼") jack (mono)

	0210400
	37204
	Paediatric rectal/oesophageal.
	YSI 400 series.
	3
	6.35mm (¼") jack (mono)

	0210403
	37804
	Paediatric skin.
	YSI 400 series.
	3
	6.35mm (¼") jack (mono)

	0210404
	37904
	Adult skin.
	YSI 400 series.
	3
	6.35mm (¼") jack (mono)

	0210701
	37107
	Adult rectal/oesophageal.
	YSI 700 series.
	3
	6.35mm (¼") jack (stereo)

	0210704
	37907
	Adult skin.
	YSI 700 series.
	3
	6.35mm (¼") jack (stereo)

	0210411
	37104/D
	Adult rectal/oesophageal.

For use with 

Dolphin 2150 pulse oximeter.
	YSI 400 series.
	3
	Miniature jack (mono)

	0210410
	37204/D
	Paediatric rectal/oesophageal.

For use with 

Dolphin 2150 pulse oximeter.
	YSI 400 series.
	3
	Miniature jack (mono)

	0210413
	37804/D
	Paediatric skin.

For use with 

Dolphin 2150 pulse oximeter.
	YSI 400 series.
	3
	Miniature jack (mono)

	0210414
	37904/D
	Adult skin.

For use with

Dolphin 2150 pulse oximeter.
	YSI 400 series.
	3
	Miniature jack (mono)


� EMBED CorelDraw.Graphic.8  ���


























[image: image1.jpg]
VIAMED Ltd.






15 Station Road, Cross Hills, 
Keighley, West, Yorkshire BD20 7DT, United Kingdom 

Tel:    +44 (0)1535 634542 Fax:   +44 (0)1535 635582

Email: info@viamed.co.uk    www.viamed-online.com
Series 400 & 700 Temperature Probes Instructions For Use

[image: image5.jpg]


[image: image6.jpg]


_1250407480.unknown

