[image: image1.png]


 

                   O2 Sensors

	COMPANY OPERATING PROCEDURES


Teledyne T7 (0110057)
O2 Sensor Test Procedure
VM3/COP/37.52
	Date:
  3-Feb-03
	Revision date: 10-Jul-20
	Issue:   3



Test & Equipment Required
Functional test required using T7 tester and cable, (TED200 oxygen analyser and cable).

Output specification as in document: (VM3COP37.99, column (i)). 

Method.
1. Open the packaging and remove the O2 sensor.

2. Check the O2 sensor for damage and signs of leaking electrolyte.

3. Connect the O2 sensor to the T7 tester

4. Check for an output Ensure that the T7 tester has been calibrated using a calibrated oxycal. If not, carry out a calibration using a calibrated oxycal.

5. If no, or very low output, or sensor is returned as faulty leave the O2 sensor exposed to air for 1 Hours.

6. Ensure that meter reading is within the output specification range, presently 16 - 26% displayed oxygen. If the O2 sensor fails the output specification test, it should left unsealed and boxed, placed in a ducket and put on the goods in desk with a short note explaining why it failed, ready to be booked in with a SRN

7. Disconnect the O2 sensor from the T7 tester.

8. Label with Viamed stickers if appropriate.

9. Reseal in serial numbered packet, integrity seal the edge of the packet with initialed and dated Viamed sticker and re-box.

VM3COP37.52 T-7 Oxygen Sensor 0110057.doc _ID422.doc     10/07/2020                          Page 1 of 1

[image: image1.png]