[image: image1.png]


 

                   O2 Sensors

	COMPANY OPERATING PROCEDURES


Teledyne R10DS (PN 0110211) & R10DV (PN 0110212)
O2 Sensor Test Procedure

VM3/COP/37.32
	Date:
  3-Feb-03
	Revision date: 28-Jul-06
	Issue:   3



Test & Equipment Required

Visual check : Digital Voltmeter 200mV setting.

Output specification as in document: (VM3COP37.00, column (i)).

Method.

1. Open the packaging and remove the O2 sensor.

2. Check the O2 sensor for damage and signs of leaking electrolyte. 

3. Check output   25mV +/- 2mV    

4. Label with Vandagraph stickers if appropriate.

5. Reseal in serial numbered packet, integrity seal the edge of the packet with initialed and dated Viamed sticker and re-box.
















VM3COP 37.32 Teledyne-R10DS-DV.doc       28/07/2006                 Page 1 of 1

[image: image1.png]