[image: image1.png]


 

                   O2 Sensors

	COMPANY OPERATING PROCEDURES


Teledyne R15 (0110015)
O2 Sensor Test Procedure

VM3/COP/37.34
	Date:
  3-Feb-03
	Revision date: 28-Jul-06
	Issue:   3



Test & Equipment Required

Functional test required using R15 / R23 tester and sensor holder, (2 x digital multimeters & adapter box).

Output specification as in document: (VM3COP37.00, column (i)). Test box converts current output to 17mV ± 3mV

Method.

1. Open the packaging and remove the O2 sensor.

2. Leaving shorting gauze and clip in place check the O2 sensor for damage and signs of leaking electrolyte.

3. Remove shorting gauze and clip and check rear of the O2 sensor, the copper tracks and rear membrane for damage or signs of leaking electrolyte.

4. Place the O2 sensor in the R15 / R23 tester sensor holder, screw on the cap and wait for both meter readings to stabilize (may take up to 1 minute)

5. Check for an output using R15/R23 tester

6. If no, or very low output, or sensor is returned as faulty leave the O2 sensor exposed to air with shorting clip in place for 1 Hour Minimum.

7. Ensure that the meter reading lies within the output specification range, 17mV ± 3mV. 
 If the O2 sensor fails the output specification test,  box unsealed , place in a docket and put on the goods in desk with a short note explaining why it failed, ready to be booked in with SRN

8. Remove the O2 sensor from the sensor holder, refit shorting gauze and clip and label with Viamed stickers if appropriate.

9. Reseal in serial numbered packet, re-box 















VM3COP 37.34 Teledyne-R15.doc    28/07/2006        Page 1 of 1

[image: image1.png]