[image: image1.jpg]ARAB HEALTH
2002

Summary of products

Pulse Oximetry Sensors Viamed  Dolphin  Aristo
Finger sensors v v v
Multi-site 'Y’ sensors v v v
Ear sensors v
Extension cables v v v
Adapter cables v v v
D.O.T. Sensors v
Wrap sensors v
Disposables v v ~

Pulse oximetry sensor repair services
Pulse oximeters - Stand alone, hand held, compact PC
Pulse oximetry simulators and testers

Oxygen sensors
Oxygen monitors and analysers

Oxygen hoods
Patient monitors
Infant resuscitation cabinets

Infant resuscitators
Radiant warmers

Phototherapy masks
Phototherapy light shields

ECG cables

Temperature probes

Supra-maximal nerve stimulator

Chart & recorder paper [
VIARMED

Imail infeid




[image: image2.png]



                                                                    CE File : Microstim.


Risk Analysis Report.

The bed is always part of the Resuscitation cabinet  and is mounted on the drop down door.

The  main risks are 

1. The patient dropping off. This should never happen as the patient should never be left unattended

2. Cross infection if the bed is damaged or nit cleaned correctly.

3. Damage to the bed if it is stowed incorrectly and collides with the equipment mounted inside the cabinet

4. The rotating bed adds other risks

1. The bed can rotate and could be left in an intermediate position between locks.

2. It could be damaged by been stored in this intermediate position

3. The bed rotates but design restrictions mean gaps in the outer wall of the remain and the height of the walls is lower that the flat fixed bed. 

4. There is a risk of the operator jamming fingers in the rotation mechanism

5. There is an increased risk of cross infection if the bed is not cleaned correctly/

1. There is provision for the removal and complete dismantling of the bed.












J.S.Lamb  05/14/08
F:\Drive S 080318\CE Files Live\Archive Microstim 080312\Microstim WP\Microstim 1998-200 Word Pro\E. Risk Analysis Report_Microstim_SW.doc                                                                              14/05/2008

Page 1

[image: image1.jpg][image: image2.png]